## CO2_primary_all_valid_data_X201
# CO2_DB is the CO2 mole fraction in the historical database (prior to 2017) (without corrections for X_loss and X_virial)
# CO2_mano_x2019 is the mole fraction determined for a manometric run, with corrections X_virial and X_loss: CO2_mano_x2019 = CO2_DB + X_virial + X_loss
# peakcycle as the time variable for calculating CO2 loss: X_loss = -1*slope*(peakcycle+2) before 2009, or X_loss = -1*slope*(peakcycle+4) from 2009
# For measurements made after 2016, the virial correction is no longer needed, and the loss correction was applied at the time of measurement
# episode_unc components include: episode_std_err, uncertainty associated with the volume ratio, average correction uncertainty during the episode; typical 0.04 ppm, 0.013%, 0.03 ppm, respectively
# for espisodes with < 3 runs, we set the std dev to the average for all primaries during that episode. This is to avoid a few unusually small std dev.
# episode_unc = sqrt((episode_std_dev/sqrt(episode_N))^2+(episode_mean*uVR/880)^2+(AVG_corr_unc)^2), where uVR = 0.119
Cylinder# | Year | Month | Day | CO2_DB | slope_(ppm/cycle) | peakcycle | X_virial | X_loss | CO2_mano_x2019 | CO2_correction_unc | xdate | Flag | xdate_mean | Episode_mean | Episode_stddev | Episode_N | AVG_corr_unc | Episode_unc
AL47-101 | 1996 | 11 | 14 | 395.99 | -0.0121 | 14 | 0.027 | 0.194 | 396.211 | 0.038 | 1996.87 | 0 | 1996.94 | 396.373 | 0.106 | 5 | 0.028241 | 0.076927
AL47-101 | 1996 | 12 | 13 | 396.28 | -0.00501 | 20 | 0.027 | 0.11 | 396.417 | 0.021 | 1996.95 | 0
AL47-101 | 1996 | 12 | 14 | 396.27 | -0.0049 | 19 | 0.027 | 0.103 | 396.4 | 0.02 | 1996.96 | 0
AL47-101 | 1996 | 12 | 15 | 396.25 | -0.01211 | 16 | 0.027 | 0.218 | 396.495 | 0.04 | 1996.96 | 0
AL47-101 | 1996 | 12 | 15 | 396.2 | -0.00594 | 17 | 0.027 | 0.113 | 396.34 | 0.022 | 1996.96 | 0
AL47-101 | 1998 | 8 | 26 | 396.67 | -0.0087 | 17 | 0.027 | 0.165 | 396.862 | 0.031 | 1998.65 | 0 | 1998.68 | 396.681 | 0.15 | 5 | 0.040131 | 0.094805
AL47-101 | 1998 | 8 | 26 | 396.38 | -0.0135 | 14 | 0.027 | 0.216 | 396.623 | 0.042 | 1998.65 | 0
AL47-101 | 1998 | 8 | 27 | 396.48 | -0.01196 | 14 | 0.027 | 0.191 | 396.698 | 0.038 | 1998.66 | 0
AL47-101 | 1998 | 9 | 24 | 396.53 | -0.012 | 15 | 0.027 | 0.204 | 396.761 | 0.039 | 1998.73 | 0
AL47-101 | 1998 | 9 | 24 | 396.31 | -0.00785 | 14 | 0.027 | 0.126 | 396.463 | 0.051 | 1998.73 | 0
AL47-101 | 2000 | 1 | 12 | 396.37 | -0.00385 | 29 | 0.033 | 0.119 | 396.523 | 0.02 | 2000.03 | 0 | 2000.04 | 396.521 | 0.044 | 3 | 0.017389 | 0.061829
AL47-101 | 2000 | 1 | 13 | 396.46 | -0.00335 | 19 | 0.033 | 0.07 | 396.563 | 0.015 | 2000.04 | 0
AL47-101 | 2000 | 1 | 13 | 396.36 | -0.00396 | 19 | 0.033 | 0.083 | 396.476 | 0.017 | 2000.04 | 0
AL47-101 | 2001 | 5 | 23 | 396.36 | -0.00456 | 35 | 0.033 | 0.169 | 396.562 | 0.025 | 2001.4 | 0 | 2001.47 | 396.513 | 0.075 | 2 | 0.02361 | 0.079025
AL47-101 | 2001 | 5 | 24 | 396.28 | -0.0040479 | 35 | 0.033 | 0.14977 | 396.463 | 0.0225 | 2001.4 | 0
AL47-101 | 2002 | 12 | 6 | 396.31 | -0.0048 | 14 | 0.033 | 0.077 | 396.42 | 0.018 | 2002.93 | 0 | 2002.99 | 396.348 | 0.071 | 4 | 0.019511 | 0.067183
AL47-101 | 2003 | 1 | 2 | 396.17 | -0.00733 | 13 | 0.033 | 0.11 | 396.313 | 0.025 | 2003.01 | 0
AL47-101 | 2003 | 1 | 4 | 396.17 | -0.0042 | 13 | 0.033 | 0.063 | 396.266 | 0.016 | 2003.01 | 0
AL47-101 | 2003 | 1 | 7 | 396.27 | -0.0047 | 17 | 0.033 | 0.089 | 396.392 | 0.019 | 2003.02 | 0
AL47-101 | 2004 | 10 | 20 | 396.34 | -0.0065 | 18 | 0.037 | 0.13 | 396.507 | 0.025 | 2004.8 | 0 | 2004.81 | 396.523 | 0.026 | 3 | 0.027342 | 0.062033
AL47-101 | 2004 | 10 | 21 | 396.31 | -0.0079 | 24 | 0.037 | 0.205 | 396.553 | 0.032 | 2004.81 | 0
AL47-101 | 2004 | 10 | 22 | 396.37 | -0.0079 | 11 | 0.037 | 0.103 | 396.51 | 0.025 | 2004.81 | 0
AL47-101 | 2006 | 12 | 11 | 396.33 | -0.0046 | 13 | 0.037 | 0.069 | 396.436 | 0.017 | 2006.95 | 0 | 2006.95 | 396.409 | 0.038891 | 2 | 0.014555 | 0.061981
AL47-101 | 2006 | 12 | 12 | 396.31 | -0.0026 | 11 | 0.037 | 0.034 | 396.381 | 0.012 | 2006.95 | 0
AL47-101 | 2009 | 10 | 8 | 396.42 | -0.00393 | 33 | 0.037 | 0.145 | 396.603 | 0.027 | 2009.77 | 0 | 2009.78 | 396.575 | 0.042 | 3 | 0.027259 | 0.064861
AL47-101 | 2009 | 10 | 9 | 396.32 | -0.00369 | 42 | 0.037 | 0.17 | 396.527 | 0.028 | 2009.77 | 0
AL47-101 | 2009 | 10 | 17 | 396.39 | -0.00351 | 44 | 0.037 | 0.168 | 396.596 | 0.027 | 2009.8 | 0
AL47-101 | 2012 | 12 | 7 | 396.3 | -0.0029 | 39 | 0.037 | 0.125 | 396.462 | 0.022 | 2012.94 | 0 | 2012.95 | 396.504 | 0.059 | 4 | 0.024382 | 0.065876
AL47-101 | 2012 | 12 | 10 | 396.39 | -0.0042 | 31 | 0.037 | 0.147 | 396.574 | 0.028 | 2012.94 | 0
AL47-101 | 2012 | 12 | 11 | 396.37 | -0.0035 | 31 | 0.037 | 0.123 | 396.53 | 0.024 | 2012.95 | 0
AL47-101 | 2012 | 12 | 13 | 396.28 | -0.0032 | 37 | 0.037 | 0.131 | 396.449 | 0.024 | 2012.95 | 0
AL47-101 | 2015 | 6 | 22 | 396.236 | -0.0029 | 57 | 0.037 | 0.177 | 396.45 | 0.093 | 2015.477 | 0 | 2015.48 | 396.486 | 0.045 | 3 | 0.026985 | 0.065405
AL47-101 | 2015 | 6 | 23 | 396.239 | -0.0029 | 63 | 0.037 | 0.194 | 396.471 | 0.093 | 2015.48 | 0
AL47-101 | 2015 | 6 | 26 | 396.314 | -0.0028 | 62 | 0.037 | 0.185 | 396.536 | 0.093 | 2015.488 | 0
AL47-102 | 1996 | 12 | 8 | 304.42 | -0.0055 | 23 | 0.016 | 0.137 | 304.573 | 0.023 | 1996.94 | 0 | 1997.11 | 304.543 | 0.084 | 2 | 0.022944 | 0.075832
AL47-102 | 1996 | 12 | 11 | 304.37 | -0.00532 | 22 | 0.016 | 0.128 | 304.514 | 0.022 | 1996.95 | 0
AL47-102 | 1998 | 8 | 31 | 304.31 | -0.0061114 | 15 | 0.016 | 0.104 | 304.43 | 0.031 | 1998.67 | 0 | 1998.68 | 304.594 | 0.136 | 5 | 0.031493 | 0.079922
AL47-102 | 1998 | 9 | 3 | 304.5 | -0.006429 | 15 | 0.016 | 0.109 | 304.625 | 0.031 | 1998.67 | 0
AL47-102 | 1998 | 9 | 4 | 304.53 | -0.0066746 | 27 | 0.016 | 0.194 | 304.74 | 0.029 | 1998.68 | 0
AL47-102 | 1998 | 9 | 10 | 304.58 | -0.0055 | 16 | 0.016 | 0.099 | 304.695 | 0.026 | 1998.69 | 0
AL47-102 | 1998 | 9 | 10 | 304.32 | -0.0069488 | 18 | 0.016 | 0.139 | 304.475 | 0.039 | 1998.69 | 0
AL47-102 | 2000 | 2 | 22 | 304.44 | -0.0024 | 30 | 0.02 | 0.077 | 304.537 | 0.014 | 2000.14 | 0 | 2000.15 | 304.511 | 0.101 | 2 | 0.014789 | 0.083755
AL47-102 | 2000 | 2 | 23 | 304.39 | -0.00311 | 22 | 0.02 | 0.075 | 304.485 | 0.015 | 2000.15 | 0
AL47-102 | 2001 | 8 | 14 | 304.42 | -0.00172 | 21 | 0.02 | 0.04 | 304.48 | 0.011 | 2001.62 | 0 | 2001.62 | 304.451 | 0.04 | 3 | 0.0093611 | 0.048124
AL47-102 | 2001 | 8 | 15 | 304.42 | -0.00132 | 19 | 0.02 | 0.028 | 304.468 | 0.009 | 2001.62 | 0
AL47-102 | 2001 | 8 | 16 | 304.36 | -0.00112 | 20 | 0.02 | 0.025 | 304.405 | 0.009 | 2001.63 | 0
AL47-102 | 2003 | 1 | 16 | 304.3 | -0.004 | 15 | 0.02 | 0.068 | 304.388 | 0.016 | 2003.04 | 0 | 2003.36 | 304.423 | 0.056 | 8 | 0.018662 | 0.049345
AL47-102 | 2003 | 1 | 16 | 304.41 | -0.0038 | 16 | 0.02 | 0.068 | 304.498 | 0.016 | 2003.04 | 0
AL47-102 | 2003 | 1 | 17 | 304.33 | -0.0033 | 17 | 0.02 | 0.063 | 304.413 | 0.015 | 2003.05 | 0
AL47-102 | 2003 | 4 | 6 | 304.25 | -0.005 | 17 | 0.02 | 0.095 | 304.365 | 0.02 | 2003.27 | 0
AL47-102 | 2003 | 4 | 6 | 304.26 | -0.0039 | 17 | 0.02 | 0.074 | 304.354 | 0.016 | 2003.27 | 0
AL47-102 | 2003 | 7 | 28 | 304.31 | -0.0044755 | 11 | 0.02 | 0.058 | 304.388 | 0.024 | 2003.58 | 0
AL47-102 | 2003 | 8 | 2 | 304.36 | -0.00679 | 11 | 0.02 | 0.088 | 304.468 | 0.023 | 2003.59 | 0
AL47-102 | 2003 | 10 | 12 | 304.39 | -0.00543 | 13 | 0.02 | 0.081 | 304.492 | 0.02 | 2003.78 | 0
AL47-102 | 2007 | 1 | 4 | 304.25 | -0.0046 | 17 | 0.022 | 0.087 | 304.36 | 0.018 | 2007.01 | 0 | 2007.01 | 304.36 | 0.061 | 1 | 0.018498 | 0.075876
AL47-102 | 2009 | 10 | 1 | 304.36 | -0.00257 | 37 | 0.022 | 0.105 | 304.488 | 0.02 | 2009.75 | 0 | 2009.75 | 304.488 | 0.077 | 1 | 0.019938 | 0.089565
AL47-102 | 2012 | 7 | 11 | 304.42 | -0.0036 | 40 | 0.022 | 0.158 | 304.601 | 0.027 | 2012.53 | 0 | 2012.66 | 304.537 | 0.082 | 5 | 0.023123 | 0.059795
AL47-102 | 2012 | 7 | 12 | 304.44 | -0.0032 | 40 | 0.022 | 0.141 | 304.603 | 0.024 | 2012.53 | 0
AL47-102 | 2012 | 7 | 27 | 304.43 | -0.003 | 39 | 0.022 | 0.129 | 304.581 | 0.023 | 2012.57 | 0
AL47-102 | 2013 | 3 | 14 | 304.29 | -0.003 | 33 | 0.022 | 0.111 | 304.423 | 0.022 | 2013.21 | 0
AL47-102 | 2013 | 3 | 15 | 304.35 | -0.0026 | 37 | 0.022 | 0.107 | 304.479 | 0.02 | 2013.21 | 0
AL47-102 | 2015 | 7 | 1 | 304.339 | -0.002 | 61 | 0.022 | 0.13 | 304.491 | 0.02 | 2015.5 | 0 | 2015.51 | 304.508 | 0.104 | 4 | 0.020234 | 0.069347
AL47-102 | 2015 | 7 | 7 | 304.517 | -0.002 | 54 | 0.022 | 0.116 | 304.655 | 0.019 | 2015.52 | 0
AL47-102 | 2015 | 7 | 8 | 304.235 | -0.0019 | 77 | 0.022 | 0.154 | 304.411 | 0.022 | 2015.52 | 0
AL47-102 | 2015 | 7 | 9 | 304.347 | -0.0023 | 41 | 0.022 | 0.104 | 304.473 | 0.019 | 2015.52 | 0
AL47-103 | 2002 | 12 | 5 | 353.15 | -0.0024 | 15 | 0.027 | 0.041 | 353.217 | 0.012 | 2002.93 | 0 | 2003.04 | 353.284 | 0.058 | 3 | 0.015 | 0.060238
AL47-103 | 2002 | 12 | 9 | 353.23 | -0.0037 | 14 | 0.027 | 0.059 | 353.316 | 0.015 | 2002.94 | 0
AL47-103 | 2003 | 4 | 2 | 353.22 | -0.0043 | 15 | 0.027 | 0.073 | 353.32 | 0.017 | 2003.26 | 0
AL47-103 | 2004 | 8 | 16 | 353.21 | -0.0057 | 16 | 0.03 | 0.103 | 353.342 | 0.021 | 2004.63 | 0 | 2004.63 | 353.389 | 0.053 | 4 | 0.037 | 0.065992
AL47-103 | 2004 | 8 | 18 | 353.17 | -0.0062 | 22 | 0.03 | 0.149 | 353.349 | 0.05 | 2004.63 | 0
AL47-103 | 2004 | 8 | 18 | 353.25 | -0.00578 | 21 | 0.03 | 0.133 | 353.413 | 0.056 | 2004.63 | 0
AL47-103 | 2004 | 8 | 20 | 353.31 | -0.0039 | 27 | 0.03 | 0.113 | 353.453 | 0.019 | 2004.64 | 0
AL47-103 | 2006 | 12 | 1 | 353.31 | -0.0052 | 11 | 0.03 | 0.068 | 353.408 | 0.018 | 2006.92 | 0 | 2006.94 | 353.343 | 0.057 | 3 | 0.013 | 0.059457
AL47-103 | 2006 | 12 | 13 | 353.2 | -0.0051 | 12 | 0.03 | 0.071 | 353.301 | 0.018 | 2006.95 | 0
AL47-103 | 2006 | 12 | 14 | 353.23 | -0.0034 | 16 | 0.03 | 0.061 | 353.321 | 0.015 | 2006.96 | 0
AL47-103 | 2009 | 9 | 1 | 353.17 | -0.0022 | 40 | 0.03 | 0.097 | 353.297 | 0.018 | 2009.67 | 0 | 2009.68 | 353.321 | 0.077 | 2 | 0.017 | 0.074406
AL47-103 | 2009 | 9 | 2 | 353.18 | -0.00315 | 39 | 0.03 | 0.135 | 353.345 | 0.024 | 2009.67 | 0
AL47-105 | 1996 | 9 | 18 | 369.41 | -0.0054 | 28 | 0.024 | 0.162 | 369.596 | 0.025 | 1996.72 | 0 | 1996.87 | 369.54 | 0.087 | 4 | 0.027819 | 0.071857
AL47-105 | 1996 | 11 | 14 | 369.27 | -0.0089 | 13 | 0.024 | 0.134 | 369.427 | 0.029 | 1996.87 | 0
AL47-105 | 1996 | 11 | 14 | 369.41 | -0.0116 | 14 | 0.024 | 0.186 | 369.619 | 0.037 | 1996.87 | 0
AL47-105 | 1997 | 1 | 6 | 369.4 | -0.0053 | 16 | 0.024 | 0.095 | 369.519 | 0.02 | 1997.02 | 0
AL47-105 | 1998 | 8 | 11 | 369.35 | -0.01075 | 10 | 0.024 | 0.129 | 369.503 | 0.032 | 1998.61 | 0 | 1998.77 | 369.52 | 0.166 | 3 | 0.043996 | 0.1167
AL47-105 | 1998 | 8 | 11 | 369.15 | -0.0075297 | 17 | 0.023 | 0.143 | 369.317 | 0.062 | 1998.61 | 0
AL47-105 | 1998 | 8 | 28 | 369.47 | -0.00948 | 24 | 0.024 | 0.247 | 369.74 | 0.038 | 1998.66 | 0
AL47-105 | 2000 | 1 | 10 | 369.24 | -0.00352 | 18 | 0.029 | 0.07 | 369.339 | 0.016 | 2000.03 | 0 | 2000.08 | 369.502 | 0.156 | 5 | 0.016731 | 0.087429
AL47-105 | 2000 | 1 | 10 | 369.2 | -0.00397 | 22 | 0.029 | 0.095 | 369.324 | 0.018 | 2000.03 | 0
AL47-105 | 2000 | 2 | 8 | 369.51 | -0.00326 | 18 | 0.029 | 0.065 | 369.604 | 0.015 | 2000.11 | 0
AL47-105 | 2000 | 2 | 9 | 369.52 | -0.00353 | 18 | 0.029 | 0.071 | 369.62 | 0.016 | 2000.11 | 0
AL47-105 | 2000 | 2 | 16 | 369.49 | -0.00469 | 20 | 0.029 | 0.103 | 369.622 | 0.02 | 2000.13 | 0
AL47-105 | 2001 | 4 | 23 | 369.56 | -0.00329 | 19 | 0.029 | 0.069 | 369.658 | 0.015 | 2001.31 | 0 | 2001.35 | 369.509 | 0.138 | 4 | 0.015035 | 0.086509
AL47-105 | 2001 | 4 | 24 | 369.45 | -0.0039 | 28 | 0.029 | 0.117 | 369.596 | 0.02 | 2001.32 | 0
AL47-105 | 2001 | 5 | 20 | 369.28 | -0.00233 | 33 | 0.029 | 0.082 | 369.391 | 0.015 | 2001.39 | 0
AL47-105 | 2001 | 5 | 21 | 369.32 | -0.00172 | 23 | 0.029 | 0.043 | 369.392 | 0.011 | 2001.39 | 0
AL47-105 | 2002 | 12 | 3 | 369.38 | -0.0039 | 17 | 0.029 | 0.074 | 369.483 | 0.016 | 2002.92 | 0 | 2003.05 | 369.474 | 0.026 | 3 | 0.017796 | 0.055121
AL47-105 | 2003 | 1 | 31 | 369.38 | -0.005 | 15 | 0.029 | 0.085 | 369.494 | 0.019 | 2003.08 | 0
AL47-105 | 2003 | 4 | 3 | 369.33 | -0.0043 | 18 | 0.029 | 0.086 | 369.445 | 0.018 | 2003.26 | 0
AL47-105 | 2004 | 8 | 9 | 369.35 | -0.0048485 | 11 | 0.033 | 0.063 | 369.446 | 0.043 | 2004.61 | 0 | 2004.61 | 369.496 | 0.044 | 3 | 0.043223 | 0.070783
AL47-105 | 2004 | 8 | 9 | 369.41 | -0.0047624 | 13 | 0.033 | 0.071 | 369.514 | 0.048 | 2004.61 | 0
AL47-105 | 2004 | 8 | 10 | 369.41 | -0.0064685 | 11 | 0.033 | 0.084 | 369.527 | 0.039 | 2004.61 | 0
AL47-105 | 2006 | 10 | 16 | 369.38 | -0.0055 | 17 | 0.033 | 0.105 | 369.517 | 0.021 | 2006.79 | 0 | 2006.8 | 369.537 | 0.061 | 2 | 0.021872 | 0.069542
AL47-105 | 2006 | 10 | 17 | 369.41 | -0.006 | 17 | 0.033 | 0.114 | 369.557 | 0.023 | 2006.8 | 0
AL47-105 | 2009 | 9 | 4 | 369.49 | -0.00373 | 43 | 0.033 | 0.175 | 369.698 | 0.028 | 2009.68 | 0 | 2009.69 | 369.572 | 0.115 | 3 | 0.025183 | 0.086834
AL47-105 | 2009 | 9 | 8 | 369.32 | -0.00242 | 46 | 0.033 | 0.121 | 369.474 | 0.021 | 2009.69 | 0
AL47-105 | 2009 | 9 | 9 | 369.33 | -0.00283 | 60 | 0.033 | 0.181 | 369.544 | 0.027 | 2009.69 | 0
AL47-105 | 2012 | 7 | 3 | 369.328 | -0.0028 | 61 | 0.033 | 0.182 | 369.543 | 0.027 | 2012.51 | 0 | 2012.51 | 369.573 | 0.078 | 3 | 0.026726 | 0.072387
AL47-105 | 2012 | 7 | 5 | 369.301 | -0.0027 | 63 | 0.033 | 0.181 | 369.514 | 0.026 | 2012.51 | 0
AL47-105 | 2012 | 7 | 6 | 369.445 | -0.00311 | 55 | 0.033 | 0.183 | 369.661 | 0.027 | 2012.52 | 0
AL47-105 | 2015 | 4 | 27 | 369.316 | -0.00306 | 44 | 0.033 | 0.147 | 369.495 | 0.024 | 2015.32 | 0 | 2015.33 | 369.521 | 0.037 | 3 | 0.025234 | 0.059917
AL47-105 | 2015 | 4 | 28 | 369.304 | -0.0033 | 47 | 0.033 | 0.168 | 369.505 | 0.027 | 2015.33 | 0
AL47-105 | 2015 | 4 | 29 | 369.374 | -0.0029 | 50 | 0.033 | 0.157 | 369.563 | 0.025 | 2015.33 | 0
AL47-106 | 1996 | 12 | 7 | 412.05 | -0.0072 | 15 | 0.029 | 0.122 | 412.202 | 0.025 | 1996.94 | 0 | 1996.96 | 412.099 | 0.103 | 3 | 0.021118 | 0.084189
AL47-106 | 1996 | 12 | 7 | 411.98 | -0.0049 | 16 | 0.029 | 0.088 | 412.098 | 0.019 | 1996.94 | 0
AL47-106 | 1997 | 1 | 4 | 411.88 | -0.0049 | 16 | 0.029 | 0.088 | 411.997 | 0.019 | 1997.01 | 0
AL47-106 | 1998 | 8 | 18 | 411.87 | -0.00785 | 18 | 0.029 | 0.157 | 412.056 | 0.024 | 1998.63 | 0 | 1998.69 | 412.267 | 0.184 | 3 | 0.031915 | 0.12414
AL47-106 | 1998 | 9 | 18 | 412.15 | -0.009863 | 16 | 0.029 | 0.178 | 412.357 | 0.027 | 1998.72 | 0
AL47-106 | 1998 | 9 | 18 | 412.2 | -0.00997 | 14 | 0.029 | 0.16 | 412.389 | 0.044 | 1998.72 | 0
AL47-106 | 2000 | 1 | 14 | 412.27 | -0.004 | 18 | 0.036 | 0.08 | 412.386 | 0.017 | 2000.04 | 0 | 2000.43 | 412.325 | 0.064 | 3 | 0.016547 | 0.068906
AL47-106 | 2000 | 1 | 14 | 412.21 | -0.00384 | 20 | 0.036 | 0.084 | 412.33 | 0.017 | 2000.04 | 0
AL47-106 | 2000 | 1 | 14 | 412.15 | -0.00345 | 19 | 0.036 | 0.073 | 412.258 | 0.016 | 2000.04 | 0
AL47-106 | 2001 | 8 | 9 | 412.26 | -0.00218 | 32 |  |  |  |  | 2001.61 | 1
AL47-106 | 2001 | 8 | 9 | 412.23 | -0.00303 | 29 | 0.036 | 0.094 | 412.36 | 0.017 | 2001.61 | 0 | 2001.62 | 412.211 | 0.145 | 3 |  | 0.10058
AL47-106 | 2001 | 8 | 9 | 411.99 | -0.00204 | 20 | 0.036 | 0.045 | 412.07 | 0.011 | 2001.61 | 0
AL47-106 | 2001 | 9 | 3 | 412.1 | -0.00188 | 34 | 0.036 | 0.068 | 412.203 | 0.013 | 2001.67 | 0
AL47-106 | 2003 | 1 | 7 | 411.97 | -0.007 | 17 | 0.036 | 0.133 | 412.139 | 0.026 | 2003.02 | 0 | 2003.15 | 412.161 | 0.018 | 4 | 0.021226 | 0.060316
AL47-106 | 2003 | 3 | 5 | 412.05 | -0.0055 | 15 | 0.036 | 0.094 | 412.179 | 0.02 | 2003.18 | 0
AL47-106 | 2003 | 3 | 11 | 412.04 | -0.0047 | 15 | 0.036 | 0.08 | 412.156 | 0.018 | 2003.2 | 0
AL47-106 | 2003 | 3 | 17 | 412.01 | -0.0042 | 28 | 0.036 | 0.126 | 412.172 | 0.021 | 2003.21 | 0
AL47-106 | 2004 | 10 | 27 | 412 | -0.0082 | 17 | 0.04 | 0.156 | 412.196 | 0.029 | 2004.82 | 0 | 2004.82 | 412.206 | 0.013 | 2 | 0.029441 | 0.063705
AL47-106 | 2004 | 10 | 28 | 411.99 | -0.0074 | 23 | 0.04 | 0.185 | 412.215 | 0.03 | 2004.83 | 0
AL47-106 | 2006 | 12 | 15 | 412.09 | -0.0073 | 11 | 0.04 | 0.095 | 412.225 | 0.024 | 2006.96 | 0 | 2006.96 | 412.234 | 0.061 | 2 | 0.022867 | 0.074101
AL47-106 | 2006 | 12 | 16 | 412.1 | -0.006 | 15 | 0.04 | 0.102 | 412.242 | 0.022 | 2006.96 | 0
AL47-106 | 2009 | 10 | 6 | 412.11 | -0.00363 | 32 | 0.04 | 0.131 | 412.281 | 0.025 | 2009.77 | 0 | 2009.78 | 412.275 | 0.124 | 3 | 0.02604 | 0.094401
AL47-106 | 2009 | 10 | 7 | 412 | -0.00287 | 34 | 0.04 | 0.109 | 412.149 | 0.021 | 2009.77 | 0
AL47-106 | 2009 | 10 | 14 | 412.16 | -0.00456 | 39 | 0.04 | 0.196 | 412.396 | 0.032 | 2009.79 | 0
AL47-106 | 2012 | 12 | 17 | 412.06 | -0.0033 | 36 | 0.04 | 0.132 | 412.232 | 0.024 | 2012.96 | 0 | 2012.98 | 412.254 | 0.044 | 4 | 0.024162 | 0.064619
AL47-106 | 2012 | 12 | 19 | 412.11 | -0.0032 | 36 | 0.04 | 0.128 | 412.278 | 0.023 | 2012.97 | 0
AL47-106 | 2012 | 12 | 26 | 412.13 | -0.0033 | 36 | 0.04 | 0.132 | 412.302 | 0.024 | 2012.99 | 0
AL47-106 | 2012 | 12 | 28 | 412.02 | -0.0035 | 37 | 0.04 | 0.144 | 412.204 | 0.025 | 2012.99 | 0
AL47-106 | 2015 | 7 | 10 | 411.927 | -0.0029 | 62 | 0.04 | 0.191 | 412.159 | 0.028 | 2015.53 | 0 | 2015.75 | 412.254 | 0.098 | 7 | 0.027031 | 0.072184
AL47-106 | 2015 | 7 | 13 | 411.978 | -0.0029 | 68 | 0.04 | 0.209 | 412.227 | 0.029 | 2015.54 | 0
AL47-106 | 2015 | 7 | 14 | 411.853 | -0.0027 | 77 | 0.04 | 0.219 | 412.112 | 0.03 | 2015.54 | 0
AL47-106 | 2015 | 11 | 6 | 412.125 | -0.0037 | 33 | 0.04 | 0.137 | 412.302 | 0.025 | 2015.85 | 0
AL47-106 | 2015 | 12 | 1 | 412.242 | -0.0039 | 28 | 0.04 | 0.125 | 412.407 | 0.025 | 2015.92 | 0
AL47-106 | 2015 | 12 | 2 | 412.089 | -0.0039 | 38 | 0.04 | 0.164 | 412.293 | 0.028 | 2015.92 | 0
AL47-106 | 2015 | 12 | 7 | 412.118 | -0.0036 | 30 | 0.04 | 0.122 | 412.281 | 0.024 | 2015.94 | 0
AL47-107 | 1996 | 11 | 22 | 453.4 | -0.00751 | 17 | 0.036 | 0.143 | 453.578 | 0.027 | 1996.89 | 0 | 1996.98 | 453.374 | 0.227 | 7 | 0.026069 | 0.10863
AL47-107 | 1996 | 11 | 23 | 452.73 | -0.0124 | 13 | 0.035 | 0.186 | 452.951 | 0.038 | 1996.9 | 0
AL47-107 | 1996 | 12 | 16 | 453.22 | -0.00601 | 19 | 0.035 | 0.126 | 453.382 | 0.023 | 1996.96 | 0
AL47-107 | 1997 | 1 | 10 | 453.39 | -0.0075 | 15 | 0.036 | 0.128 | 453.553 | 0.026 | 1997.03 | 0
AL47-107 | 1997 | 1 | 10 | 453.38 | -0.0071 | 16 | 0.036 | 0.128 | 453.543 | 0.025 | 1997.03 | 0
AL47-107 | 1997 | 1 | 11 | 453.24 | -0.00613 | 18 | 0.035 | 0.123 | 453.398 | 0.023 | 1997.03 | 0
AL47-107 | 1997 | 1 | 12 | 453.09 | -0.0048 | 16 | 0.035 | 0.086 | 453.212 | 0.019 | 1997.03 | 0
AL47-107 | 1998 | 8 | 19 | 453.28 | -0.0070686 | 15 | 0.035 | 0.12 | 453.436 | 0.055 | 1998.64 | 0 | 1998.81 | 453.41 | 0.057 | 3 | 0.040691 | 0.080611
AL47-107 | 1998 | 8 | 19 | 453.27 | -0.009 | 14 | 0.035 | 0.144 | 453.449 | 0.03 | 1998.64 | 0
AL47-107 | 1998 | 9 | 22 | 453.15 | -0.0088786 | 16 | 0.035 | 0.16 | 453.345 | 0.037 | 1998.73 | 0
AL47-107 | 2000 | 1 | 19 | 453.16 | -0.0045 | 22 | 0.043 | 0.108 | 453.311 | 0.02 | 2000.05 | 0 | 2000.07 | 453.224 | 0.153 | 5 | 0.018381 | 0.09368
AL47-107 | 2000 | 1 | 20 | 453.2 | -0.00379 | 22 | 0.043 | 0.091 | 453.334 | 0.017 | 2000.05 | 0
AL47-107 | 2000 | 1 | 20 | 453.1 | -0.00385 | 21 | 0.043 | 0.089 | 453.231 | 0.017 | 2000.05 | 0
AL47-107 | 2000 | 2 | 7 | 452.82 | -0.00462 | 19 | 0.043 | 0.097 | 452.959 | 0.019 | 2000.1 | 0
AL47-107 | 2000 | 2 | 8 | 453.15 | -0.0043 | 19 | 0.043 | 0.09 | 453.283 | 0.018 | 2000.11 | 0
AL47-107 | 2001 | 5 | 26 | 452.99 | -0.005 | 50 |  |  |  |  | 2001.4 | 1
AL47-107 | 2001 | 6 | 3 | 452.96 | -0.00389 | 22 | 0.043 | 0.093 | 453.096 | 0.018 | 2001.42 | 0 | 2001.42 | 453.119 | 0.075 | 2 | 0.020417 | 0.08357
AL47-107 | 2001 | 6 | 4 | 452.94 | -0.00387 | 39 | 0.043 | 0.159 | 453.141 | 0.023 | 2001.43 | 0
AL47-107 | 2002 | 11 | 29 | 453.03 | -0.0074 | 14 | 0.043 | 0.118 | 453.191 | 0.025 | 2002.91 | 0 | 2003.01 | 453.152 | 0.07 | 3 | 0.027009 | 0.078217
AL47-107 | 2002 | 12 | 31 | 453 | -0.0101 | 13 | 0.043 | 0.152 | 453.194 | 0.032 | 2003 | 0
AL47-107 | 2003 | 2 | 14 | 452.91 | -0.0063 | 17 | 0.043 | 0.12 | 453.072 | 0.023 | 2003.12 | 0
AL47-107 | 2004 | 10 | 29 | 453 | -0.0083 | 12 | 0.049 | 0.116 | 453.165 | 0.027 | 2004.83 | 0 | 2004.84 | 453.206 | 0.041 | 3 | 0.029171 | 0.071884
AL47-107 | 2004 | 11 | 2 | 453.02 | -0.0087 | 14 | 0.049 | 0.139 | 453.208 | 0.029 | 2004.84 | 0
AL47-107 | 2004 | 11 | 3 | 453.03 | -0.0093 | 16 | 0.049 | 0.167 | 453.246 | 0.032 | 2004.84 | 0
AL47-107 | 2006 | 11 | 14 | 453.08 | -0.0065 | 14 | 0.049 | 0.104 | 453.233 | 0.023 | 2006.87 | 0 | 2006.92 | 453.181 | 0.075 | 3 | 0.021909 | 0.07817
AL47-107 | 2006 | 11 | 15 | 453.07 | -0.006 | 14 | 0.049 | 0.096 | 453.215 | 0.021 | 2006.87 | 0
AL47-107 | 2007 | 1 | 10 | 452.96 | -0.0062 | 12 | 0.049 | 0.087 | 453.095 | 0.021 | 2007.03 | 0
AL47-107 | 2009 | 9 | 18 | 453.11 | -0.00427 | 42 | 0.049 | 0.196 | 453.355 | 0.031 | 2009.72 | 0 | 2009.72 | 453.285 | 0.099 | 2 | 0.028673 | 0.097365
AL47-107 | 2009 | 9 | 22 | 453.02 | -0.00365 | 36 | 0.049 | 0.146 | 453.215 | 0.026 | 2009.73 | 0
AL47-107 | 2013 | 1 | 9 | 453.1 | -0.0042 | 32 | 0.049 | 0.151 | 453.3 | 0.028 | 2013.02 | 0 | 2013.03 | 453.273 | 0.046 | 3 | 0.026852 | 0.071996
AL47-107 | 2013 | 1 | 11 | 453.02 | -0.0037 | 37 | 0.049 | 0.152 | 453.22 | 0.027 | 2013.03 | 0
AL47-107 | 2013 | 1 | 15 | 453.11 | -0.0038 | 33 | 0.049 | 0.141 | 453.299 | 0.026 | 2013.04 | 0
AL47-107 | 2015 | 7 | 20 | 452.93 | -0.0033 | 60 | 0.049 | 0.211 | 453.19 | 0.03 | 2015.55 | 0 | 2015.69 | 453.203 | 0.04 | 6 | 0.02997 | 0.070148
AL47-107 | 2015 | 7 | 21 | 452.9 | -0.0034 | 52 | 0.049 | 0.19 | 453.139 | 0.029 | 2015.56 | 0
AL47-107 | 2015 | 7 | 22 | 452.92 | -0.0031 | 69 | 0.049 | 0.226 | 453.195 | 0.031 | 2015.56 | 0
AL47-107 | 2015 | 10 | 22 | 452.938 | -0.0033 | 65 | 0.049 | 0.228 | 453.214 | 0.032 | 2015.81 | 0
AL47-107 | 2015 | 10 | 28 | 453.029 | -0.0037 | 45 | 0.049 | 0.181 | 453.259 | 0.029 | 2015.83 | 0
AL47-107 | 2015 | 10 | 30 | 452.975 | -0.0037 | 49 | 0.049 | 0.196 | 453.22 | 0.03 | 2015.83 | 0
AL47-110 | 1996 | 11 | 16 | 246.56 | -0.0073 | 12 | 0.011 | 0.102 | 246.673 | 0.024 | 1996.88 | 0 | 1996.94 | 246.673 | 0.084 | 1 | 0.024259 | 0.09358
AL47-110 | 1998 | 9 | 14 | 246.6 | -0.0050903 | 16 | 0.011 | 0.092 | 246.702 | 0.012 | 1998.71 | 0 | 1998.71 | 246.696 | 0.035 | 3 | 0.034094 | 0.051804
AL47-110 | 1998 | 9 | 15 | 246.63 | -0.005808 | 13 | 0.011 | 0.087 | 246.728 | 0.051 | 1998.71 | 0
AL47-110 | 1998 | 9 | 16 | 246.56 | -0.0055 | 14 | 0.011 | 0.088 | 246.659 | 0.039 | 1998.71 | 0
AL47-110 | 1999 | 12 | 8 | 246.83 | -0.00195 | 21 | 0.011 | 0.045 | 246.888 | 0.011 | 1999.94 | 0 | 1999.9833 | 246.8617 | 0.027025 | 3 | 0.034 | 0.050138
AL47-110 | 1999 | 12 | 9 | 246.8 | -0.00216 | 21 | 0.011 | 0.05 | 246.863 | 0.012 | 1999.94 | 0
AL47-110 | 2000 | 1 | 26 | 246.78 | -0.00201 | 18 | 0.013 | 0.04 | 246.834 | 0.011 | 2000.07 | 0
AL47-110 | 2001 | 8 | 18 | 246.69 | -0.00142 | 28 | 0.013 | 0.043 | 246.746 | 0.01 | 2001.63 | 0 | 2001.64 | 246.783 | 0.032 | 3 | 0.009935 | 0.039417
AL47-110 | 2001 | 8 | 19 | 246.75 | -0.0016 | 21 | 0.013 | 0.037 | 246.8 | 0.01 | 2001.64 | 0
AL47-110 | 2001 | 8 | 20 | 246.76 | -0.00145 | 18 | 0.013 | 0.029 | 246.802 | 0.009 | 2001.64 | 0
AL47-110 | 2003 | 1 | 23 | 246.62 | -0.0037 | 16 | 0.013 | 0.067 | 246.7 | 0.016 | 2003.06 | 0 | 2003.07 | 246.699 | 0.01 | 3 | 0.013622 | 0.036494
AL47-110 | 2003 | 1 | 24 | 246.65 | -0.0021 | 19 | 0.013 | 0.044 | 246.708 | 0.011 | 2003.07 | 0
AL47-110 | 2003 | 1 | 27 | 246.62 | -0.003 | 17 | 0.013 | 0.057 | 246.69 | 0.014 | 2003.07 | 0
AL47-110 | 2004 | 11 | 4 | 246.65 | -0.004 | 14 | 0.015 | 0.064 | 246.729 | 0.016 | 2004.84 | 0 | 2004.85 | 246.755 | 0.043 | 2 | 0.017478 | 0.048409
AL47-110 | 2004 | 11 | 5 | 246.7 | -0.0055 | 10 | 0.015 | 0.066 | 246.781 | 0.019 | 2004.85 | 0
AL47-110 | 2006 | 12 | 20 | 246.53 | -0.0028 | 15 | 0.015 | 0.048 | 246.593 | 0.013 | 2006.97 | 0 | 2006.99 | 246.676 | 0.074 | 3 | 0.013639 | 0.055893
AL47-110 | 2006 | 12 | 28 | 246.67 | -0.0025 | 18 | 0.015 | 0.05 | 246.735 | 0.013 | 2006.99 | 0
AL47-110 | 2007 | 1 | 8 | 246.63 | -0.004 | 12 | 0.015 | 0.056 | 246.701 | 0.016 | 2007.02 | 0
AL47-110 | 2010 | 2 | 3 | 246.53 | -0.00277 | 21 | 0.015 | 0.069 | 246.614 | 0.018 | 2010.09 | 0 | 2010.11 | 246.65 | 0.056 | 3 | 0.017404 | 0.049606
AL47-110 | 2010 | 2 | 11 | 246.53 | -0.0023 | 29 | 0.015 | 0.076 | 246.621 | 0.017 | 2010.11 | 0
AL47-110 | 2010 | 2 | 12 | 246.63 | -0.00227 | 27 | 0.015 | 0.07 | 246.715 | 0.017 | 2010.12 | 0
AL47-110 | 2010 | 2 | 18 | 246.68 | -0.00234 | 31 | 0.015 | 0.082 | 246.777 | 0.018 | 2010.13 | 0
AL47-110 | 2012 | 8 | 21 | 246.51 | -0.0026 | 48 | 0.015 | 0.135 | 246.66 | 0.022 | 2012.64 | 0 | 2012.65 | 246.769 | 0.113 | 4 | 0.020875 | 0.068859
AL47-110 | 2012 | 8 | 22 | 246.56 | -0.0022 | 48 | 0.015 | 0.114 | 246.689 | 0.02 | 2012.64 | 0
AL47-110 | 2012 | 8 | 24 | 246.75 | -0.0022 | 55 | 0.015 | 0.13 | 246.895 | 0.021 | 2012.65 | 0
AL47-110 | 2012 | 8 | 27 | 246.69 | -0.0023 | 51 | 0.015 | 0.127 | 246.832 | 0.021 | 2012.66 | 0
AL47-110 | 2015 | 5 | 29 | 246.629 | -0.0018 | 42 | 0.015 | 0.083 | 246.727 | 0.016 | 2015.41 | 0 | 2015.42 | 246.692 | 0.034 | 4 | 0.017298 | 0.041244
AL47-110 | 2015 | 6 | 1 | 246.531 | -0.002 | 46 | 0.015 | 0.1 | 246.646 | 0.018 | 2015.42 | 0
AL47-110 | 2015 | 6 | 2 | 246.578 | -0.0017 | 60 | 0.015 | 0.109 | 246.702 | 0.018 | 2015.42 | 0
AL47-110 | 2015 | 6 | 3 | 246.587 | -0.0019 | 45 | 0.015 | 0.093 | 246.695 | 0.017 | 2015.42 | 0
AL47-111 | 1996 | 8 | 29 | 323.82 | -0.003 | 28 | 0.018 | 0.09 | 323.928 | 0.016 | 1996.66 | 0 | 1996.75 | 324.077 | 0.146 | 5 | 0.016736 | 0.080398
AL47-111 | 1996 | 9 | 3 | 323.84 | -0.0032 | 18 | 0.018 | 0.064 | 323.922 | 0.015 | 1996.67 | 0
AL47-111 | 1996 | 9 | 5 | 324.12 | -0.0045 | 23 | 0.018 | 0.113 | 324.251 | 0.02 | 1996.68 | 0
AL47-111 | 1996 | 9 | 9 | 324.05 | -0.0023 | 25 | 0.018 | 0.062 | 324.13 | 0.013 | 1996.69 | 0
AL47-111 | 1997 | 1 | 16 | 324.04 | -0.0049 | 18 | 0.018 | 0.098 | 324.156 | 0.02 | 1997.04 | 0
AL47-111 | 1998 | 9 | 2 | 323.85 | -0.0057352 | 21 | 0.018 | 0.132 | 324 | 0.042 | 1998.67 | 0 | 1998.69 | 324.017 | 0.069 | 4 | 0.038917 | 0.068005
AL47-111 | 1998 | 9 | 2 | 323.97 | -0.0056923 | 20 | 0.018 | 0.125 | 324.113 | 0.043 | 1998.67 | 0
AL47-111 | 1998 | 9 | 10 | 323.78 | -0.00788 | 17 | 0.018 | 0.15 | 323.948 | 0.028 | 1998.69 | 0
AL47-111 | 1998 | 9 | 11 | 323.89 | -0.0068995 | 12 | 0.018 | 0.097 | 324.005 | 0.043 | 1998.7 | 0
AL47-111 | 2000 | 1 | 4 | 324.11 | -0.00261 | 20 | 0.023 | 0.057 | 324.19 | 0.013 | 2000.01 | 0 | 2000.05 | 324.243 | 0.127 | 4 | 0.013832 | 0.078397
AL47-111 | 2000 | 1 | 5 | 324.28 | -0.0027 | 18 | 0.023 | 0.054 | 324.357 | 0.013 | 2000.01 | 0
AL47-111 | 2000 | 1 | 5 | 324.25 | -0.00313 | 18 | 0.023 | 0.063 | 324.335 | 0.014 | 2000.01 | 0
AL47-111 | 2000 | 2 | 24 | 323.99 | -0.0027 | 26 | 0.023 | 0.075 | 324.088 | 0.015 | 2000.15 | 0
AL47-111 | 2001 | 6 | 5 | 323.97 | -0.00311 | 41 | 0.023 | 0.134 | 324.126 | 0.02 | 2001.43 | 0 | 2001.59 | 324.126 | 0.099 | 5 | 0.015829 | 0.06428
AL47-111 | 2001 | 8 | 16 | 323.93 | -0.00193 | 50 | 0.023 | 0.1 | 324.053 | 0.016 | 2001.63 | 0
AL47-111 | 2001 | 8 | 17 | 324.19 | -0.00181 | 44 | 0.023 | 0.083 | 324.296 | 0.014 | 2001.63 | 0
AL47-111 | 2001 | 8 | 17 | 323.98 | -0.00182 | 43 | 0.023 | 0.082 | 324.085 | 0.014 | 2001.63 | 0
AL47-111 | 2001 | 8 | 17 | 323.96 | -0.00181 | 47 | 0.023 | 0.089 | 324.071 | 0.015 | 2001.63 | 0
AL47-111 | 2003 | 1 | 31 | 324.01 | -0.0034 | 16 | 0.023 | 0.061 | 324.094 | 0.015 | 2003.08 | 0 | 2003.17 | 324.068 | 0.079 | 4 | 0.014663 | 0.060792
AL47-111 | 2003 | 2 | 21 | 324.08 | -0.0038 | 15 | 0.023 | 0.065 | 324.167 | 0.016 | 2003.14 | 0
AL47-111 | 2003 | 3 | 6 | 323.93 | -0.0021 | 17 | 0.023 | 0.04 | 323.992 | 0.011 | 2003.18 | 0
AL47-111 | 2003 | 4 | 3 | 323.92 | -0.0042 | 16 | 0.023 | 0.076 | 324.018 | 0.017 | 2003.26 | 0
AL47-111 | 2004 | 10 | 8 | 324 | -0.0054 | 17 | 0.025 | 0.103 | 324.128 | 0.021 | 2004.77 | 0 | 2004.78 | 324.112 | 0.042 | 3 | 0.021674 | 0.054578
AL47-111 | 2004 | 10 | 11 | 324.01 | -0.0054 | 18 | 0.025 | 0.108 | 324.143 | 0.021 | 2004.78 | 0
AL47-111 | 2004 | 10 | 12 | 323.93 | -0.0064 | 15 | 0.025 | 0.109 | 324.064 | 0.023 | 2004.78 | 0
AL47-111 | 2006 | 10 | 31 | 324.09 | -0.0064 | 14 | 0.025 | 0.102 | 324.218 | 0.023 | 2006.83 | 0 | 2006.85 | 324.147 | 0.1 | 2 | 0.021497 | 0.085927
AL47-111 | 2006 | 11 | 9 | 323.94 | -0.0048 | 21 | 0.025 | 0.11 | 324.076 | 0.02 | 2006.86 | 0
AL47-111 | 2009 | 9 | 23 | 324.02 | -0.00272 | 42 | 0.025 | 0.125 | 324.17 | 0.022 | 2009.73 | 0 | 2009.73 | 324.139 | 0.077 | 2 | 0.020552 | 0.072857
AL47-111 | 2009 | 9 | 24 | 323.98 | -0.00243 | 38 | 0.025 | 0.102 | 324.107 | 0.019 | 2009.73 | 0
AL47-111 | 2012 | 8 | 28 | 324.04 | -0.0036 | 49 | 0.025 | 0.191 | 324.256 | 0.029 | 2012.66 | 0 | 2012.81 | 324.228 | 0.115 | 6 | 0.025884 | 0.069256
AL47-111 | 2012 | 8 | 31 | 324.13 | -0.0036 | 45 | 0.025 | 0.176 | 324.332 | 0.028 | 2012.67 | 0
AL47-111 | 2012 | 9 | 7 | 324.14 | -0.0033 | 45 | 0.025 | 0.162 | 324.327 | 0.026 | 2012.69 | 0
AL47-111 | 2012 | 9 | 19 | 324.07 | -0.0039 | 45 | 0.025 | 0.191 | 324.286 | 0.03 | 2012.72 | 0
AL47-111 | 2013 | 2 | 13 | 323.95 | -0.0029 | 31 | 0.025 | 0.102 | 324.077 | 0.021 | 2013.12 | 0
AL47-111 | 2013 | 2 | 15 | 323.95 | -0.0028 | 37 | 0.025 | 0.115 | 324.09 | 0.021 | 2013.12 | 0
AL47-111 | 2015 | 5 | 6 | 324.022 | -0.0027 | 39 | 0.025 | 0.116 | 324.163 | 0.021 | 2015.35 | 0 | 2015.36 | 324.158 | 0.101 | 3 | 0.021535 | 0.076063
AL47-111 | 2015 | 5 | 8 | 324.104 | -0.0025 | 47 | 0.025 | 0.128 | 324.257 | 0.021 | 2015.36 | 0
AL47-111 | 2015 | 5 | 12 | 323.887 | -0.0023 | 58 | 0.025 | 0.143 | 324.055 | 0.022 | 2015.37 | 0
AL47-121 | 1996 | 9 | 8 | 349.21 | -0.0068 | 21 | 0.021 | 0.156 | 349.387 | 0.027 | 1996.69 | 0 | 1996.88 | 349.474 | 0.066 | 4 | 0.023903 | 0.062399
AL47-121 | 1996 | 9 | 10 | 349.31 | -0.00655 | 21 | 0.021 | 0.151 | 349.482 | 0.026 | 1996.69 | 0
AL47-121 | 1997 | 1 | 4 | 349.37 | -0.0053 | 15 | 0.021 | 0.09 | 349.481 | 0.02 | 1997.01 | 0
AL47-121 | 1997 | 1 | 5 | 349.41 | -0.0061 | 17 | 0.021 | 0.116 | 349.547 | 0.023 | 1997.01 | 0
AL47-121 | 1998 | 8 | 13 | 349.37 | -0.00888 | 14 | 0.021 | 0.142 | 349.533 | 0.038 | 1998.62 | 0 | 1998.64 | 349.487 | 0.033 | 3 | 0.028117 | 0.058199
AL47-121 | 1998 | 8 | 14 | 349.35 | -0.0072 | 16 | 0.021 | 0.13 | 349.501 | 0.026 | 1998.62 | 0
AL47-121 | 1998 | 9 | 9 | 349.36 | -0.00575 | 15 | 0.021 | 0.098 | 349.479 | 0.021 | 1998.69 | 0
AL47-121 | 2000 | 1 | 28 | 349.45 | -0.00306 | 19 | 0.026 | 0.064 | 349.54 | 0.014 | 2000.08 | 0 | 2000.08 | 349.56 | 0.054 | 3 | 0.014533 | 0.058461
AL47-121 | 2000 | 1 | 29 | 349.53 | -0.0031 | 19 | 0.026 | 0.065 | 349.621 | 0.015 | 2000.08 | 0
AL47-121 | 2000 | 1 | 31 | 349.43 | -0.0033 | 17 | 0.026 | 0.063 | 349.519 | 0.015 | 2000.08 | 0
AL47-121 | 2001 | 5 | 15 | 349.38 | -0.00286 | 33 | 0.026 | 0.1 | 349.506 | 0.017 | 2001.37 | 0 | 2001.39 | 349.53 | 0.046 | 3 | 0.017959 | 0.057113
AL47-121 | 2001 | 5 | 15 | 349.43 | -0.00318 | 38 | 0.026 | 0.127 | 349.583 | 0.02 | 2001.37 | 0
AL47-121 | 2001 | 6 | 5 | 349.37 | -0.003 | 33 | 0.026 | 0.105 | 349.501 | 0.017 | 2001.43 | 0
AL47-121 | 2002 | 12 | 4 | 349.45 | -0.0035 | 16 | 0.026 | 0.063 | 349.539 | 0.015 | 2002.93 | 0 | 2002.94 | 349.465 | 0.064 | 3 | 0.016399 | 0.062189
AL47-121 | 2002 | 12 | 10 | 349.34 | -0.0044 | 14 | 0.026 | 0.07 | 349.436 | 0.017 | 2002.94 | 0
AL47-121 | 2002 | 12 | 16 | 349.35 | -0.005 | 7 | 0.026 | 0.045 | 349.421 | 0.017 | 2002.96 | 0
AL47-121 | 2006 | 11 | 7 | 349.37 | -0.0053 | 20 | 0.029 | 0.117 | 349.516 | 0.022 | 2006.85 | 0 | 2006.86 | 349.504 | 0.043 | 2 | 0.020324 | 0.05976
AL47-121 | 2006 | 11 | 8 | 349.38 | -0.0051 | 14 | 0.029 | 0.082 | 349.491 | 0.019 | 2006.86 | 0
AL47-121 | 2009 | 9 | 10 | 349.38 | -0.00244 | 39 | 0.029 | 0.105 | 349.514 | 0.02 | 2009.69 | 0 | 2009.73 | 349.514 | 0.077 | 1 | 0.019517 | 0.092433
AL47-121 | 2012 | 10 | 12 | 349.48 | -0.0036 | 40 | 0.029 | 0.158 | 349.668 | 0.027 | 2012.78 | 0 | 2012.82 | 349.663 | 0.07 | 4 | 0.028365 | 0.065309
AL47-121 | 2012 | 10 | 16 | 349.4 | -0.0031 | 39 | 0.029 | 0.133 | 349.563 | 0.023 | 2012.79 | 0
AL47-121 | 2012 | 11 | 6 | 349.47 | -0.005 | 40 | 0.029 | 0.22 | 349.719 | 0.035 | 2012.85 | 0
AL47-121 | 2012 | 11 | 9 | 349.5 | -0.0037 | 43 | 0.029 | 0.174 | 349.703 | 0.028 | 2012.86 | 0
AL47-121 | 2015 | 6 | 12 | 349.236 | -0.0023 | 70 | 0.029 | 0.17 | 349.435 | 0.024 | 2015.45 | 0 | 2015.46 | 349.434 | 0.073 | 5 | 0.02227 | 0.0616
AL47-121 | 2015 | 6 | 16 | 349.155 | -0.0024 | 54 | 0.029 | 0.139 | 349.323 | 0.022 | 2015.46 | 0
AL47-121 | 2015 | 6 | 17 | 349.321 | -0.0025 | 57 | 0.029 | 0.153 | 349.503 | 0.023 | 2015.46 | 0
AL47-121 | 2015 | 6 | 18 | 349.239 | -0.0019 | 71 | 0.029 | 0.143 | 349.411 | 0.021 | 2015.47 | 0
AL47-121 | 2015 | 6 | 19 | 349.36 | -0.0026 | 37 | 0.029 | 0.107 | 349.496 | 0.02 | 2015.47 | 0
AL47-123 | 1996 | 11 | 20 | 422.8 | -0.0129 | 13 | 0.031 | 0.193 | 423.024 | 0.04 | 1996.89 | 0 | 1997.08 | 423.132 | 0.073 | 4 | 0.030626 | 0.074459
AL47-123 | 1996 | 11 | 24 | 422.99 | -0.00923 | 15 | 0.031 | 0.157 | 423.178 | 0.031 | 1996.9 | 0
AL47-123 | 1997 | 1 | 6 | 423.03 | -0.0054 | 15 | 0.031 | 0.092 | 423.153 | 0.02 | 1997.02 | 0
AL47-123 | 1997 | 1 | 6 | 422.98 | -0.0095 | 15 | 0.031 | 0.162 | 423.172 | 0.032 | 1997.02 | 0
AL47-123 | 1998 | 8 | 17 | 422.86 | -0.0075363 | 14 | 0.031 | 0.121 | 423.011 | 0.07 | 1998.63 | 0 | 1998.82 | 423.249 | 0.217 | 3 | 0.048081 | 0.14589
AL47-123 | 1998 | 9 | 20 | 423.15 | -0.010075 | 10 | 0.031 | 0.121 | 423.302 | 0.037 | 1998.72 | 0
AL47-123 | 1998 | 9 | 21 | 423.22 | -0.0115 | 14 | 0.031 | 0.184 | 423.435 | 0.037 | 1998.72 | 0
AL47-123 | 2000 | 1 | 17 | 423.09 | -0.00321 | 24 | 0.037 | 0.083 | 423.211 | 0.016 | 2000.05 | 0 | 2000.06 | 423.35 | 0.196 | 2 | 0.018139 | 0.15104
AL47-123 | 2000 | 1 | 18 | 423.33 | -0.00417 | 27 | 0.037 | 0.121 | 423.488 | 0.02 | 2000.05 | 0
AL47-123 | 2001 | 6 | 7 | 422.96 | -0.0037 | 37 | 0.037 | 0.144 | 423.142 | 0.022 | 2001.44 | 0 | 2001.56 | 423.198 | 0.113 | 4 | 0.023391 | 0.083752
AL47-123 | 2001 | 8 | 7 | 423.11 | -0.00381 | 54 | 0.037 | 0.213 | 423.361 | 0.028 | 2001.6 | 0
AL47-123 | 2001 | 8 | 8 | 422.9 | -0.00302 | 54 | 0.037 | 0.169 | 423.107 | 0.023 | 2001.61 | 0
AL47-123 | 2001 | 8 | 8 | 423 | -0.00249 | 57 | 0.037 | 0.147 | 423.184 | 0.021 | 2001.61 | 0
AL47-123 | 2003 | 2 | 1 | 423.09 | -0.0062 | 15 | 0.037 | 0.105 | 423.233 | 0.022 | 2003.09 | 0 | 2003.16 | 423.163 | 0.086 | 6 | 0.020502 | 0.070196
AL47-123 | 2003 | 2 | 19 | 422.89 | -0.0048 | 16 | 0.037 | 0.086 | 423.014 | 0.019 | 2003.14 | 0
AL47-123 | 2003 | 2 | 20 | 423.1 | -0.0053 | 15 | 0.037 | 0.09 | 423.228 | 0.02 | 2003.14 | 0
AL47-123 | 2003 | 2 | 20 | 422.99 | -0.0052 | 15 | 0.037 | 0.088 | 423.116 | 0.02 | 2003.14 | 0
AL47-123 | 2003 | 3 | 10 | 423.07 | -0.0065 | 15 | 0.037 | 0.111 | 423.218 | 0.023 | 2003.19 | 0
AL47-123 | 2003 | 4 | 1 | 423.04 | -0.0049 | 16 | 0.037 | 0.088 | 423.166 | 0.019 | 2003.25 | 0
AL47-123 | 2004 | 10 | 13 | 423.18 | -0.0089 | 15 | 0.042 | 0.151 | 423.374 | 0.03 | 2004.79 | 0 | 2004.79 | 423.314 | 0.043 | 4 | 0.027033 | 0.066857
AL47-123 | 2004 | 10 | 14 | 423.12 | -0.0077 | 14 | 0.042 | 0.123 | 423.286 | 0.026 | 2004.79 | 0
AL47-123 | 2004 | 10 | 14 | 423.13 | -0.0076 | 12 | 0.042 | 0.106 | 423.279 | 0.025 | 2004.79 | 0
AL47-123 | 2004 | 10 | 15 | 423.14 | -0.0078 | 15 | 0.042 | 0.133 | 423.315 | 0.027 | 2004.79 | 0
AL47-123 | 2006 | 10 | 26 | 423.14 | -0.0073 | 19 | 0.042 | 0.153 | 423.336 | 0.027 | 2006.82 | 0 | 2006.82 | 423.336 | 0.061 | 1 | 0.027393 | 0.088026
AL47-123 | 2009 | 9 | 1 | 422.95 | -0.00308 | 43 | 0.042 | 0.145 | 423.137 | 0.024 | 2009.67 | 0 | 2009.7 | 423.221 | 0.118 | 2 | 0.025693 | 0.10439
AL47-123 | 2009 | 9 | 25 | 423.1 | -0.00367 | 40 | 0.042 | 0.161 | 423.304 | 0.027 | 2009.74 | 0
AL47-123 | 2013 | 1 | 3 | 423.1 | -0.00333 | 36 | 0.042 | 0.133 | 423.276 | 0.024 | 2013.01 | 0 | 2013.01 | 423.214 | 0.058 | 3 | 0.023699 | 0.070415
AL47-123 | 2013 | 1 | 4 | 423.03 | -0.0032 | 37 | 0.042 | 0.131 | 423.204 | 0.024 | 2013.01 | 0
AL47-123 | 2013 | 1 | 8 | 423 | -0.0034 | 31 | 0.042 | 0.119 | 423.161 | 0.023 | 2013.02 | 0
AL47-123 | 2015 | 7 | 15 | 422.988 | -0.0032 | 56 | 0.042 | 0.192 | 423.222 | 0.028 | 2015.54 | 0 | 2015.54 | 423.19 | 0.028 | 3 | 0.02796 | 0.065711
AL47-123 | 2015 | 7 | 16 | 422.922 | -0.0029 | 67 | 0.042 | 0.206 | 423.17 | 0.029 | 2015.54 | 0
AL47-123 | 2015 | 7 | 17 | 422.97 | -0.0034 | 45 | 0.042 | 0.167 | 423.179 | 0.027 | 2015.55 | 0
AL47-130 | 1996 | 11 | 16 | 337.14 | -0.0064 | 9 | 0.02 | 0.07 | 337.23 | 0.021 | 1996.88 | 0 | 1996.92 | 337.27 | 0.084 | 2 | 0.01865 | 0.077175
AL47-130 | 1997 | 1 | 3 | 337.2 | -0.0031 | 27 | 0.02 | 0.09 | 337.31 | 0.016 | 1997.01 | 0
AL47-130 | 1998 | 9 | 1 | 337.51 | -0.0060086 | 13 | 0.02 | 0.09 | 337.62 | 0.052 | 1998.67 | 0 | 1998.68 | 337.488 | 0.114 | 3 | 0.050619 | 0.094747
AL47-130 | 1998 | 9 | 1 | 337.28 | -0.00778 | 15 | 0.02 | 0.132 | 337.432 | 0.042 | 1998.67 | 0
AL47-130 | 1998 | 9 | 7 | 337.27 | -0.007255 | 15 | 0.02 | 0.123 | 337.413 | 0.057 | 1998.69 | 0
AL47-130 | 2000 | 1 | 6 | 337.24 | -0.00321 | 18 | 0.024 | 0.064 | 337.329 | 0.015 | 2000.02 | 0 | 2000.02 | 337.37 | 0.036 | 3 | 0.016223 | 0.052693
AL47-130 | 2000 | 1 | 6 | 337.3 | -0.00377 | 17 | 0.024 | 0.072 | 337.396 | 0.016 | 2000.02 | 0
AL47-130 | 2000 | 1 | 7 | 337.26 | -0.0037 | 25 | 0.024 | 0.1 | 337.384 | 0.018 | 2000.02 | 0
AL47-130 | 2001 | 5 | 5 | 337.25 | -0.00377 | 41 | 0.024 | 0.162 | 337.437 | 0.023 | 2001.35 | 0 | 2001.37 | 337.397 | 0.075 | 2 | 0.019749 | 0.072692
AL47-130 | 2001 | 5 | 17 | 337.23 | -0.00205 | 48 | 0.024 | 0.102 | 337.357 | 0.016 | 2001.38 | 0
AL47-130 | 2001 | 5 | 22 | 337.26 | -0.00275 | 23 | 0.024 | 0.069 | 337.353 | 0.014 | 2001.39 | 0
AL47-130 | 2002 | 12 | 18 | 337.29 | -0.0057 | 13 | 0.024 | 0.086 | 337.4 | 0.02 | 2002.97 | 0 | 2003.04 | 337.325 | 0.072 | 5 | 0.017864 | 0.058623
AL47-130 | 2002 | 12 | 18 | 337.17 | -0.0033 | 14 | 0.024 | 0.053 | 337.247 | 0.014 | 2002.97 | 0
AL47-130 | 2002 | 12 | 19 | 337.3 | -0.0045 | 16 | 0.024 | 0.081 | 337.405 | 0.018 | 2002.97 | 0
AL47-130 | 2003 | 1 | 8 | 337.17 | -0.0048 | 16 | 0.024 | 0.086 | 337.281 | 0.019 | 2003.02 | 0
AL47-130 | 2003 | 4 | 7 | 337.19 | -0.0047 | 15 | 0.024 | 0.08 | 337.294 | 0.018 | 2003.27 | 0
AL47-130 | 2004 | 10 | 7 | 337.3 | -0.0047 | 15 | 0.027 | 0.08 | 337.407 | 0.018 | 2004.77 | 0 | 2004.77 | 337.453 | 0.048 | 3 | 0.021474 | 0.057546
AL47-130 | 2004 | 10 | 7 | 337.32 | -0.006 | 15 | 0.027 | 0.102 | 337.449 | 0.022 | 2004.77 | 0
AL47-130 | 2004 | 10 | 8 | 337.35 | -0.0066 | 17 | 0.027 | 0.125 | 337.503 | 0.024 | 2004.77 | 0
AL47-130 | 2006 | 10 | 24 | 337.23 | -0.0047 | 14 | 0.027 | 0.075 | 337.333 | 0.018 | 2006.82 | 0 | 2006.83 | 337.363 | 0.043 | 2 | 0.018064 | 0.057724
AL47-130 | 2006 | 10 | 25 | 337.28 | -0.0045 | 17 | 0.027 | 0.086 | 337.393 | 0.018 | 2006.82 | 0
AL47-130 | 2009 | 11 | 2 | 337.29 | -0.00349 | 34 | 0.027 | 0.133 | 337.45 | 0.025 | 2009.84 | 0 | 2009.86 | 337.415 | 0.046 | 3 | 0.020504 | 0.056636
AL47-130 | 2009 | 11 | 5 | 337.24 | -0.00251 | 34 | 0.027 | 0.095 | 337.363 | 0.019 | 2009.85 | 0
AL47-130 | 2009 | 11 | 19 | 337.32 | -0.00235 | 32 | 0.027 | 0.085 | 337.432 | 0.018 | 2009.89 | 0
AL47-130 | 2012 | 9 | 24 | 337.34 | -0.0031 | 49 | 0.027 | 0.164 | 337.532 | 0.026 | 2012.73 | 0 | 2012.75 | 337.51 | 0.055 | 3 | 0.025254 | 0.061067
AL47-130 | 2012 | 9 | 25 | 337.26 | -0.0029 | 51 | 0.027 | 0.16 | 337.447 | 0.025 | 2012.74 | 0
AL47-130 | 2012 | 10 | 4 | 337.38 | -0.0034 | 38 | 0.027 | 0.143 | 337.55 | 0.025 | 2012.76 | 0
AL47-130 | 2015 | 6 | 5 | 337.333 | -0.0027 | 45 | 0.027 | 0.132 | 337.493 | 0.022 | 2015.43 | 0 | 2015.44 | 337.436 | 0.055 | 4 | 0.022125 | 0.057688
AL47-130 | 2015 | 6 | 9 | 337.191 | -0.0021 | 64 | 0.027 | 0.143 | 337.361 | 0.022 | 2015.44 | 0
AL47-130 | 2015 | 6 | 10 | 337.283 | -0.0026 | 48 | 0.027 | 0.135 | 337.446 | 0.022 | 2015.44 | 0
AL47-130 | 2015 | 6 | 11 | 337.279 | -0.0026 | 48 | 0.027 | 0.135 | 337.442 | 0.022 | 2015.45 | 0
AL47-132 | 1997 | 1 | 13 | 521.28 | -0.00723 | 18 | 0.047 | 0.145 | 521.472 | 0.027 | 1997.04 | 0 | 1997.04 | 521.448 | 0.056 | 5 | 0.026236 | 0.079295
AL47-132 | 1997 | 1 | 13 | 521.31 | -0.00559 | 23 | 0.047 | 0.14 | 521.497 | 0.024 | 1997.04 | 0
AL47-132 | 1997 | 1 | 14 | 521.2 | -0.00668 | 19 | 0.047 | 0.14 | 521.387 | 0.026 | 1997.04 | 0
AL47-132 | 1997 | 1 | 14 | 521.33 | -0.0062 | 17 | 0.047 | 0.118 | 521.495 | 0.023 | 1997.04 | 0
AL47-132 | 1997 | 1 | 15 | 521.16 | -0.00901 | 18 | 0.047 | 0.18 | 521.387 | 0.032 | 1997.04 | 0
AL47-132 | 1998 | 8 | 19 | 521.26 | -0.010239 | 10 | 0.047 | 0.123 | 521.43 | 0.068 | 1998.64 | 0 | 1998.7 | 521.601 | 0.21 | 5 | 0.055486 | 0.1299
AL47-132 | 1998 | 8 | 21 | 521.4 | -0.0136 | 18 | 0.047 | 0.272 | 521.719 | 0.067 | 1998.64 | 0
AL47-132 | 1998 | 9 | 25 | 521.38 | -0.0119 | 14 | 0.047 | 0.19 | 521.617 | 0.029 | 1998.74 | 0
AL47-132 | 1998 | 9 | 25 | 52.08 | -0.013811 | 15 | 0.047 | 0.234 | 521.362 | 0.031 | 1998.74 | 0
AL47-132 | 1998 | 9 | 30 | 521.64 | -0.011857 | 14 | 0.047 | 0.19 | 521.877 | 0.057 | 1998.75 | 0
AL47-132 | 1999 | 12 | 7 | 521.31 | -0.00407 | 20 | 0.047 | 0.089 | 521.455 | 0.018 | 1999.94 | 0 | 1999.94 | 521.601 | 0.218 | 2 | 0.018567 | 0.17053
AL47-132 | 1999 | 12 | 8 | 521.6 | -0.00419 | 24 | 0.047 | 0.109 | 521.765 | 0.019 | 1999.94 | 0
AL47-132 | 2001 | 8 | 2 | 521.51 | -0.00455 | 25 | 0.056 | 0.123 | 521.688 | 0.021 | 2001.59 | 0 | 2001.62 | 521.644 | 0.125 | 6 | 0.021905 | 0.089777
AL47-132 | 2001 | 8 | 3 | 521.49 | -0.00354 | 56 | 0.056 | 0.205 | 521.751 | 0.027 | 2001.59 | 0
AL47-132 | 2001 | 8 | 6 | 521.42 | -0.00526 | 27 | 0.056 | 0.153 | 521.628 | 0.024 | 2001.6 | 0
AL47-132 | 2001 | 8 | 20 | 521.21 | -0.00295 | 45 | 0.055 | 0.138 | 521.404 | 0.02 | 2001.64 | 0
AL47-132 | 2001 | 8 | 21 | 521.47 | -0.003 | 47 | 0.056 | 0.147 | 521.672 | 0.021 | 2001.64 | 0
AL47-132 | 2001 | 8 | 21 | 521.55 | -0.00236 | 47 | 0.056 | 0.116 | 521.721 | 0.018 | 2001.64 | 0
AL47-132 | 2003 | 2 | 24 | 521.46 | -0.0063 | 15 | 0.056 | 0.107 | 521.623 | 0.023 | 2003.15 | 0 | 2003.19 | 521.627 | 0.077 | 4 | 0.024007 | 0.08387
AL47-132 | 2003 | 2 | 24 | 521.52 | -0.0075 | 15 | 0.056 | 0.128 | 521.703 | 0.026 | 2003.15 | 0
AL47-132 | 2003 | 3 | 1 | 521.49 | -0.006 | 17 | 0.056 | 0.114 | 521.66 | 0.023 | 2003.17 | 0
AL47-132 | 2003 | 3 | 30 | 521.34 | -0.0067 | 17 | 0.056 | 0.127 | 521.523 | 0.025 | 2003.25 | 0
AL47-132 | 2004 | 10 | 22 | 521.44 | -0.0107 | 12 | 0.064 | 0.15 | 521.654 | 0.033 | 2004.81 | 0 | 2004.82 | 521.663 | 0.054 | 3 | 0.03499 | 0.084692
AL47-132 | 2004 | 10 | 25 | 521.36 | -0.01 | 17 | 0.064 | 0.19 | 521.614 | 0.034 | 2004.82 | 0
AL47-132 | 2004 | 10 | 26 | 521.44 | -0.0108 | 18 | 0.064 | 0.216 | 521.72 | 0.037 | 2004.82 | 0
AL47-132 | 2006 | 11 | 21 | 521.4 | -0.0046 | 12 | 0.064 | 0.064 | 521.528 | 0.017 | 2006.89 | 0 | 2006.93 | 521.581 | 0.082 | 5 | 0.020636 | 0.082131
AL47-132 | 2006 | 11 | 22 | 521.34 | -0.0055 | 14 | 0.064 | 0.088 | 521.492 | 0.02 | 2006.89 | 0
AL47-132 | 2006 | 11 | 27 | 521.42 | -0.0076 | 15 | 0.064 | 0.129 | 521.613 | 0.026 | 2006.91 | 0
AL47-132 | 2006 | 11 | 28 | 521.51 | -0.0068 | 17 | 0.064 | 0.129 | 521.703 | 0.025 | 2006.91 | 0
AL47-132 | 2007 | 1 | 11 | 521.45 | -0.0035 | 14 | 0.064 | 0.056 | 521.57 | 0.015 | 2007.03 | 0
AL47-132 | 2009 | 10 | 22 | 521.52 | -0.00524 | 34 | 0.064 | 0.199 | 521.783 | 0.034 | 2009.81 | 0 | 2009.84 | 521.67 | 0.105 | 3 | 0.02733 | 0.096945
AL47-132 | 2009 | 11 | 6 | 521.46 | -0.00307 | 38 | 0.064 | 0.129 | 521.653 | 0.023 | 2009.85 | 0
AL47-132 | 2009 | 11 | 13 | 521.37 | -0.00328 | 39 | 0.064 | 0.141 | 521.575 | 0.025 | 2009.87 | 0
AL47-132 | 2013 | 1 | 25 | 521.32 | -0.003 | 39 | 0.064 | 0.129 | 521.513 | 0.023 | 2013.07 | 0 | 2013.09 | 521.599 | 0.065 | 4 | 0.026929 | 0.082198
AL47-132 | 2013 | 1 | 31 | 521.37 | -0.004 | 36 | 0.064 | 0.16 | 521.594 | 0.028 | 2013.08 | 0
AL47-132 | 2013 | 2 | 7 | 521.44 | -0.00403 | 37 | 0.064 | 0.165 | 521.669 | 0.028 | 2013.1 | 0
AL47-132 | 2013 | 2 | 8 | 521.39 | -0.00403 | 37 | 0.064 | 0.165 | 521.619 | 0.028 | 2013.1 | 0
AL47-132 | 2015 | 4 | 16 | 521.407 | -0.0046 | 42 | 0.064 | 0.212 | 521.683 | 0.033 | 2015.29 | 0 | 2015.3 | 521.669 | 0.072 | 4 | 0.033532 | 0.086005
AL47-132 | 2015 | 4 | 17 | 521.318 | -0.0044 | 50 | 0.064 | 0.238 | 521.62 | 0.035 | 2015.3 | 0
AL47-132 | 2015 | 4 | 20 | 521.503 | -0.0046 | 39 | 0.064 | 0.198 | 521.765 | 0.032 | 2015.3 | 0
AL47-132 | 2015 | 4 | 21 | 521.333 | -0.0046 | 42 | 0.064 | 0.212 | 521.609 | 0.033 | 2015.31 | 0
AL47-136 | 1996 | 9 | 3 | 381.23 | -0.0041 | 24 | 0.025 | 0.107 | 381.362 | 0.019 | 1996.67 | 0 | 1996.9 | 381.401 | 0.105 | 4 | 0.026108 | 0.078089
AL47-136 | 1996 | 9 | 13 | 381.08 | -0.00498 | 33 | 0.025 | 0.174 | 381.28 | 0.026 | 1996.7 | 0
AL47-136 | 1996 | 9 | 13 | 381.31 | -0.00637 | 28 | 0.025 | 0.191 | 381.526 | 0.029 | 1996.7 | 0
AL47-136 | 1996 | 11 | 12 | 381.26 | -0.0094 | 14 | 0.025 | 0.15 | 381.435 | 0.031 | 1996.87 | 0
AL47-136 | 1998 | 8 | 12 | 381.42 | -0.0073675 | 14 | 0.025 | 0.118 | 381.563 | 0.052 | 1998.62 | 0 | 1998.75 | 381.597 | 0.121 | 4 | 0.051631 | 0.094809
AL47-136 | 1998 | 8 | 13 | 381.62 | -0.0085226 | 13 | 0.025 | 0.128 | 381.773 | 0.046 | 1998.62 | 0
AL47-136 | 1998 | 8 | 15 | 381.42 | -0.006663 | 15 | 0.025 | 0.113 | 381.558 | 0.059 | 1998.62 | 0
AL47-136 | 1998 | 8 | 28 | 381.31 | -0.0093842 | 15 | 0.025 | 0.16 | 381.495 | 0.05 | 1998.66 | 0
AL47-136 | 2000 | 1 | 21 | 381.44 | -0.00352 | 20 | 0.031 | 0.078 | 381.548 | 0.016 | 2000.06 | 0 | 2000.07 | 381.463 | 0.076 | 3 | 0.017122 | 0.069853
AL47-136 | 2000 | 1 | 24 | 381.32 | -0.00426 | 18 | 0.031 | 0.085 | 381.436 | 0.018 | 2000.07 | 0
AL47-136 | 2000 | 1 | 24 | 381.28 | -0.00373 | 23 | 0.031 | 0.093 | 381.404 | 0.018 | 2000.07 | 0
AL47-136 | 2001 | 4 | 20 | 381.39 | -0.00357 | 44 | 0.031 | 0.164 | 381.585 | 0.023 | 2001.3 | 0 | 2001.35 | 381.534 | 0.041 | 4 | 0.020286 | 0.059107
AL47-136 | 2001 | 5 | 7 | 381.37 | -0.00421 | 33 | 0.031 | 0.148 | 381.548 | 0.023 | 2001.35 | 0
AL47-136 | 2001 | 5 | 8 | 381.38 | -0.00297 | 32 | 0.031 | 0.101 | 381.512 | 0.017 | 2001.36 | 0
AL47-136 | 2001 | 5 | 23 | 381.35 | -0.00333 | 31 | 0.031 | 0.11 | 381.491 | 0.018 | 2001.4 | 0
AL47-136 | 2002 | 12 | 13 | 381.31 | -0.0054 | 14 | 0.031 | 0.086 | 381.427 | 0.02 | 2002.95 | 0 | 2003.03 | 381.402 | 0.034 | 4 | 0.017203 | 0.056965
AL47-136 | 2002 | 12 | 13 | 381.27 | -0.003 | 15 | 0.031 | 0.051 | 381.352 | 0.013 | 2002.95 | 0
AL47-136 | 2002 | 12 | 16 | 381.31 | -0.0046 | 15 | 0.031 | 0.078 | 381.419 | 0.018 | 2002.96 | 0
AL47-136 | 2003 | 4 | 2 | 381.3 | -0.0043 | 17 | 0.031 | 0.082 | 381.412 | 0.018 | 2003.26 | 0
AL47-136 | 2004 | 8 | 6 | 381.33 | -0.0064603 | 13 | 0.035 | 0.097 | 381.462 | 0.032 | 2004.6 | 0 | 2004.6 | 381.438 | 0.0648 | 3 | 0.045331 | 0.078199
AL47-136 | 2004 | 8 | 7 | 381.24 | -0.00695 | 11 | 0.035 | 0.09 | 381.365 | 0.046 | 2004.6 | 0
AL47-136 | 2004 | 8 | 8 | 381.37 | -0.0069231 | 10 | 0.035 | 0.083 | 381.488 | 0.058 | 2004.61 | 0
AL47-136 | 2006 | 10 | 11 | 381.33 | -0.0064 | 17 | 0.035 | 0.122 | 381.486 | 0.024 | 2006.78 | 0 | 2006.78 | 381.524 | 0.053 | 2 | 0.024359 | 0.068262
AL47-136 | 2006 | 10 | 12 | 381.39 | -0.0065 | 19 | 0.035 | 0.137 | 381.561 | 0.025 | 2006.78 | 0
AL47-136 | 2009 | 10 | 2 | 381.34 | -0.00216 | 32 | 0.035 | 0.078 | 381.452 | 0.017 | 2009.76 | 0 | 2009.77 | 381.445 | 0.044 | 3 | 0.020818 | 0.061151
AL47-136 | 2009 | 10 | 5 | 381.3 | -0.00329 | 42 | 0.035 | 0.151 | 381.486 | 0.025 | 2009.76 | 0
AL47-136 | 2009 | 10 | 15 | 381.26 | -0.00273 | 34 | 0.035 | 0.104 | 381.398 | 0.02 | 2009.79 | 0
AL47-136 | 2012 | 11 | 30 | 381.31 | -0.003 | 38 | 0.035 | 0.126 | 381.471 | 0.023 | 2012.92 | 0 | 2012.93 | 381.505 | 0.056 | 4 | 0.023685 | 0.063297
AL47-136 | 2012 | 12 | 4 | 381.31 | -0.003 | 31 | 0.035 | 0.105 | 381.45 | 0.021 | 2012.93 | 0
AL47-136 | 2012 | 12 | 5 | 381.37 | -0.0043 | 35 | 0.035 | 0.168 | 381.572 | 0.029 | 2012.93 | 0
AL47-136 | 2012 | 12 | 6 | 381.37 | -0.0026 | 44 | 0.035 | 0.125 | 381.529 | 0.021 | 2012.93 | 0
AL47-136 | 2015 | 5 | 1 | 381.422 | -0.0032 | 43 | 0.035 | 0.15 | 381.607 | 0.025 | 2015.34 | 0 | 2015.34 | 381.57 | 0.046 | 3 | 0.024959 | 0.063172
AL47-136 | 2015 | 5 | 4 | 381.385 | -0.003 | 51 | 0.035 | 0.165 | 381.585 | 0.026 | 2015.34 | 0
AL47-136 | 2015 | 5 | 5 | 381.329 | -0.0028 | 51 | 0.035 | 0.154 | 381.518 | 0.024 | 2015.35 | 0
AL47-139 | 1996 | 9 | 17 | 361.03 | -0.0037 | 23 | 0.022 | 0.093 | 361.145 | 0.017 | 1996.71 | 0 | 1996.76 | 361.084 | 0.044 | 4 | 0.023366 | 0.058431
AL47-139 | 1996 | 9 | 17 | 360.94 | -0.0045 | 25 | 0.022 | 0.122 | 361.084 | 0.021 | 1996.71 | 0
AL47-139 | 1996 | 9 | 18 | 360.85 | -0.0078 | 23 | 0.022 | 0.195 | 361.067 | 0.031 | 1996.72 | 0
AL47-139 | 1996 | 11 | 21 | 360.9 | -0.00666 | 16 | 0.022 | 0.12 | 361.042 | 0.024 | 1996.89 | 0
AL47-139 | 1998 | 8 | 10 | 360.89 | -0.01044 | 13 | 0.022 | 0.157 | 361.069 | 0.033 | 1998.61 | 0 | 1998.81 | 361.106 | 0.055 | 3 | 0.037037 | 0.069026
AL47-139 | 1998 | 9 | 17 | 360.92 | -0.0085929 | 14 | 0.022 | 0.137 | 361.08 | 0.044 | 1998.71 | 0
AL47-139 | 1998 | 9 | 17 | 360.97 | -0.01041 | 15 | 0.022 | 0.177 | 361.169 | 0.034 | 1998.71 | 0
AL47-139 | 1999 | 12 | 10 | 360.86 | -0.00332 | 21 | 0.022 | 0.076 | 360.964 | 0.016 | 1999.94 | 0 | 2000 | 361.011 | 0.101 | 2 | 0.017279 | 0.088217
AL47-139 | 2000 | 2 | 16 | 360.92 | -0.0038 | 27 | 0.028 | 0.11 | 361.058 | 0.019 | 2000.13 | 0
AL47-139 | 2001 | 5 | 21 | 360.96 | -0.00234 | 27 | 0.028 | 0.068 | 361.056 | 0.014 | 2001.39 | 0 | 2001.39 | 361.027 | 0.047 | 3 | 0.016925 | 0.058363
AL47-139 | 2001 | 5 | 22 | 360.86 | -0.00339 | 23 | 0.028 | 0.085 | 360.973 | 0.016 | 2001.39 | 0
AL47-139 | 2001 | 5 | 22 | 360.89 | -0.00381 | 33 | 0.028 | 0.133 | 361.051 | 0.021 | 2001.39 | 0
AL47-139 | 2002 | 12 | 17 | 360.9 | -0.0043 | 15 | 0.028 | 0.073 | 361.001 | 0.017 | 2002.96 | 0 | 2003.01 | 361.006 | 0.015 | 4 | 0.016839 | 0.052182
AL47-139 | 2002 | 12 | 17 | 360.88 | -0.005 | 14 | 0.028 | 0.08 | 360.988 | 0.019 | 2002.96 | 0
AL47-139 | 2003 | 1 | 10 | 360.91 | -0.0045 | 17 | 0.028 | 0.086 | 361.023 | 0.018 | 2003.03 | 0
AL47-139 | 2003 | 1 | 30 | 360.93 | -0.0029 | 16 | 0.028 | 0.052 | 361.01 | 0.013 | 2003.08 | 0
AL47-139 | 2004 | 8 | 11 | 360.82 | -0.0051961 | 10 | 0.031 | 0.062 | 360.914 | 0.058 | 2004.61 | 0 | 2004.62 | 361.016 | 0.093 | 3 | 0.041824 | 0.083759
AL47-139 | 2004 | 8 | 12 | 360.92 | -0.0059091 | 12 | 0.031 | 0.083 | 361.034 | 0.041 | 2004.62 | 0
AL47-139 | 2004 | 8 | 13 | 360.9 | -0.006 | 26 | 0.031 | 0.168 | 361.099 | 0.026 | 2004.62 | 0
AL47-139 | 2006 | 10 | 20 | 360.94 | -0.0057 | 18 | 0.031 | 0.114 | 361.085 | 0.022 | 2006.8 | 0 | 2007.54 | 361.07 | 0.05 | 3 | 0.022192 | 0.060908
AL47-139 | 2006 | 11 | 3 | 360.92 | -0.0048 | 11 | 0.031 | 0.062 | 361.014 | 0.017 | 2006.84 | 0
AL47-139 | 2009 | 9 | 14 | 360.9 | -0.0031 | 54 | 0.031 | 0.18 | 361.111 | 0.027 | 2009.71 | 0 | 2009.71 | 361.098 | 0.077 | 2 | 0.026556 | 0.077808
AL47-139 | 2009 | 9 | 15 | 360.89 | -0.00321 | 47 | 0.031 | 0.164 | 361.085 | 0.026 | 2009.71 | 0
AL47-139 | 2012 | 11 | 15 | 360.9 | -0.0033 | 39 | 0.031 | 0.142 | 361.073 | 0.025 | 2012.87 | 0 | 2012.9 | 361.083 | 0.049 | 4 | 0.024025 | 0.05968
AL47-139 | 2012 | 11 | 27 | 360.99 | -0.0036 | 32 | 0.031 | 0.13 | 361.151 | 0.025 | 2012.91 | 0
AL47-139 | 2012 | 11 | 28 | 360.88 | -0.0035 | 31 | 0.031 | 0.123 | 361.034 | 0.024 | 2012.91 | 0
AL47-139 | 2012 | 11 | 29 | 360.92 | -0.0031 | 36 | 0.031 | 0.124 | 361.075 | 0.023 | 2012.91 | 0
AL47-139 | 2015 | 5 | 13 | 360.916 | -0.0031 | 39 | 0.031 | 0.133 | 361.08 | 0.023 | 2015.37 | 0 | 2015.38 | 361.052 | 0.028 | 3 | 0.024099 | 0.056797
AL47-139 | 2015 | 5 | 14 | 360.837 | -0.0027 | 54 | 0.031 | 0.157 | 361.025 | 0.024 | 2015.37 | 0
AL47-139 | 2015 | 5 | 27 | 360.872 | -0.0031 | 44 | 0.031 | 0.149 | 361.052 | 0.025 | 2015.41 | 0
AL47-146 | 1996 | 9 | 19 | 389.53 | -0.0053 | 21 | 0.026 | 0.122 | 389.678 | 0.022 | 1996.72 | 0 | 1997.05 | 389.61 | 0.062 | 3 | 0.021471 | 0.067217
AL47-146 | 1996 | 11 | 13 | 389.43 | -0.0066 | 13 | 0.026 | 0.099 | 389.555 | 0.023 | 1996.87 | 0
AL47-146 | 1997 | 1 | 5 | 389.48 | -0.0051 | 16 | 0.026 | 0.092 | 389.598 | 0.02 | 1997.01 | 0
AL47-146 | 1998 | 8 | 24 | 389.27 | -0.00739 | 20 |  |  |  |  | 1998.65 | 1
AL47-146 | 1998 | 9 | 23 | 389.7 | -0.0076509 | 14 | 0.026 | 0.122 | 389.849 | 0.047 | 1998.73 | 0 | 1998.86 | 389.833 | 0.121 | 2 | 0.045 | 0.11011
AL47-146 | 1998 | 9 | 23 | 389.62 | -0.0081 | 19 | 0.026 | 0.1701 | 389.816 | 0.043 | 1998.73 | 0
AL47-146 | 2000 | 1 | 11 | 389.52 | -0.00362 | 16 | 0.032 | 0.065 | 389.617 | 0.015 | 2000.03 | 0 | 2000.06 | 389.635 | 0.115 | 5 | 0.013353 | 0.07483
AL47-146 | 2000 | 1 | 12 | 389.4 | -0.00154 | 19 | 0.032 | 0.032 | 389.464 | 0.01 | 2000.03 | 0
AL47-146 | 2000 | 1 | 12 | 389.53 | -0.00291 | 19 | 0.032 | 0.061 | 389.623 | 0.014 | 2000.03 | 0
AL47-146 | 2000 | 2 | 9 | 389.59 | -0.00333 | 21 | 0.032 | 0.077 | 389.699 | 0.016 | 2000.11 | 0
AL47-146 | 2000 | 2 | 9 | 389.69 | -0.00214 | 22 | 0.032 | 0.051 | 389.773 | 0.012 | 2000.11 | 0
AL47-146 | 2001 | 5 | 31 | 389.59 | -0.00375 | 18 | 0.032 | 0.075 | 389.697 | 0.016 | 2001.42 | 0 | 2001.43 | 389.711 | 0.012 | 4 | 0.016673 | 0.055599
AL47-146 | 2001 | 6 | 1 | 389.57 | -0.00394 | 25 | 0.032 | 0.106 | 389.708 | 0.019 | 2001.42 | 0
AL47-146 | 2001 | 6 | 4 | 389.59 | -0.00314 | 27 | 0.032 | 0.091 | 389.713 | 0.017 | 2001.43 | 0
AL47-146 | 2001 | 6 | 4 | 389.61 | -0.00259 | 30 | 0.032 | 0.083 | 389.725 | 0.015 | 2001.43 | 0
AL47-146 | 2002 | 12 | 19 | 389.59 | -0.0089 | 11 | 0.032 | 0.116 | 389.738 | 0.028 | 2002.97 | 0 | 2003.09 | 389.698 | 0.059 | 5 | 0.023457 | 0.063431
AL47-146 | 2002 | 12 | 20 | 389.56 | -0.01 | 12 | 0.032 | 0.14 | 389.732 | 0.031 | 2002.97 | 0
AL47-146 | 2002 | 12 | 20 | 389.59 | -0.0086 | 12 | 0.032 | 0.12 | 389.742 | 0.028 | 2002.97 | 0
AL47-146 | 2003 | 4 | 4 | 389.49 | -0.0048 | 16 | 0.032 | 0.086 | 389.608 | 0.019 | 2003.26 | 0
AL47-146 | 2003 | 4 | 5 | 389.59 | -0.002 | 21 | 0.032 | 0.046 | 389.668 | 0.011 | 2003.26 | 0
AL47-146 | 2004 | 10 | 18 | 389.6 | -0.00736 | 21 | 0.036 | 0.169 | 389.806 | 0.029 | 2004.63 | 0 | 2004.63 | 389.788 | 0.042 | 2 | 0.026509 | 0.066054
AL47-146 | 2004 | 10 | 18 | 389.62 | -0.00705 | 14 | 0.036 | 0.113 | 389.769 | 0.024 | 2004.63 | 0
AL47-146 | 2004 | 10 | 19 | 389.64 | -0.00673 | 14 | 0.036 | 0.108 | 389.784 | 0.023 | 2004.64 | 0
AL47-146 | 2006 | 11 | 10 | 389.64 | -0.0055 | 14 | 0.036 | 0.088 | 389.764 | 0.02 | 2006.86 | 0 | 2006.86 | 389.765 | 0.061 | 2 | 0.018319 | 0.070527
AL47-146 | 2006 | 11 | 13 | 389.66 | -0.0041 | 15 | 0.036 | 0.07 | 389.766 | 0.017 | 2006.87 | 0
AL47-146 | 2009 | 8 | 19 | 389.54 | -0.0024 | 39 | 0.036 | 0.103 | 389.679 | 0.019 | 2009.64 | 0 | 2009.64 | 389.703 | 0.077 | 2 | 0.020837 | 0.078586
AL47-146 | 2009 | 8 | 21 | 389.56 | -0.00275 | 44 | 0.036 | 0.132 | 389.728 | 0.022 | 2009.64 | 0
AL47-146 | 2012 | 8 | 2 | 389.72 | -0.003 | 51 | 0.036 | 0.165 | 389.921 | 0.026 | 2012.59 | 0
AL47-146 | 2012 | 8 | 9 | 389.67 | -0.0036 | 47 | 0.036 | 0.184 | 389.89 | 0.029 | 2012.61 | 0 | 2013.24 | 389.844 | 0.078 | 5 | 0.02572 | 0.068246
AL47-146 | 2012 | 8 | 16 | 389.74 | -0.0035 | 47 | 0.036 | 0.179 | 389.955 | 0.028 | 2012.63 | 0
AL47-146 | 2012 | 8 | 20 | 389.6 | -0.003 | 47 | 0.036 | 0.153 | 389.789 | 0.025 | 2012.64 | 0
AL47-146 | 2013 | 2 | 11 | 389.61 | -0.0035 | 30 | 0.036 | 0.119 | 389.765 | 0.024 | 2013.11 | 0
AL47-146 | 2013 | 2 | 12 | 389.66 | -0.0036 | 30 | 0.036 | 0.122 | 389.819 | 0.024 | 2013.12 | 0
AL47-146 | 2015 | 4 | 22 | 389.458 | -0.0033 | 41 | 0.036 | 0.149 | 389.643 | 0.025 | 2015.31 | 0 | 2015.31 | 389.712 | 0.064 | 3 | 0.025469 | 0.069219
AL47-146 | 2015 | 4 | 23 | 389.578 | -0.0033 | 43 | 0.036 | 0.155 | 389.769 | 0.026 | 2015.31 | 0
AL47-146 | 2015 | 4 | 24 | 389.534 | -0.0033 | 43 | 0.036 | 0.155 | 389.725 | 0.026 | 2015.32 | 0
ND17440 | 2004 | 11 | 7 | 479.5 | -0.0088 | 13 | 0.054 | 0.132 | 479.686 | 0.029 | 2004.85 | 0 | 2004.86 | 479.658 | 0.03 | 3 | 0.02993 | 0.073505
ND17440 | 2004 | 11 | 9 | 479.42 | -0.0085 | 16 | 0.054 | 0.153 | 479.627 | 0.029 | 2004.86 | 0
ND17440 | 2004 | 11 | 10 | 479.44 | -0.0093 | 16 | 0.054 | 0.167 | 479.662 | 0.032 | 2004.86 | 0
ND17440 | 2007 | 1 | 22 | 479.61 | -0.0041 | 15 | 0.054 | 0.07 | 479.734 | 0.017 | 2007.06 | 0 | 2007.06 | 479.708 | 0.061 | 2 | 0.01667 | 0.079665
ND17440 | 2007 | 1 | 23 | 479.56 | -0.0043 | 14 | 0.054 | 0.069 | 479.683 | 0.017 | 2007.06 | 0
ND17440 | 2009 | 11 | 9 | 479.61 | -0.00377 | 34 | 0.054 | 0.143 | 479.808 | 0.026 | 2009.86 | 0 | 2009.86 | 479.739 | 0.061 | 3 | 0.024606 | 0.07781
ND17440 | 2009 | 11 | 10 | 479.51 | -0.0032 | 36 | 0.054 | 0.128 | 479.692 | 0.023 | 2009.86 | 0
ND17440 | 2009 | 11 | 12 | 479.53 | -0.0034 | 35 | 0.054 | 0.133 | 479.717 | 0.024 | 2009.87 | 0
ND17440 | 2013 | 1 | 22 | 479.51 | -0.0047 | 28 | 0.054 | 0.15 | 479.715 | 0.03 | 2013.06 | 0 | 2013.06 | 479.716 | 0.037 | 3 | 0.02757 | 0.073652
ND17440 | 2013 | 1 | 23 | 479.56 | -0.0042 | 29 | 0.054 | 0.139 | 479.753 | 0.027 | 2013.06 | 0
ND17440 | 2013 | 1 | 24 | 479.48 | -0.0037 | 35 | 0.054 | 0.144 | 479.679 | 0.026 | 2013.07 | 0
ND17440 | 2015 | 4 | 8 | 479.461 | -0.004 | 44 | 0.054 | 0.192 | 479.707 | 0.03 | 2015.27 | 0 | 2015.28 | 479.761 | 0.054 | 4 | 0.030119 | 0.076454
ND17440 | 2015 | 4 | 9 | 479.499 | -0.0042 | 38 | 0.054 | 0.176 | 479.73 | 0.03 | 2015.27 | 0
ND17440 | 2015 | 4 | 14 | 479.613 | -0.0042 | 34 | 0.054 | 0.16 | 479.827 | 0.029 | 2015.29 | 0
ND17440 | 2015 | 4 | 15 | 479.52 | -0.0042 | 45 | 0.054 | 0.206 | 479.78 | 0.032 | 2015.29 | 0
CC71578 | 2012 | 2 | 2 | 549.417 | -0.00364 | 32 | 0.071 | 0.131 | 549.619 | 0.025 | 2012.09 | 0 | 2012.28 | 549.612 | 0.066 | 4 | 0.029177 | 0.086395
CC71578 | 2012 | 2 | 6 | 549.473 | -0.00409 | 34 | 0.071 | 0.155 | 549.699 | 0.028 | 2012.1 | 0
CC71578 | 2012 | 2 | 7 | 549.37 | -0.00388 | 34 | 0.071 | 0.147 | 549.588 | 0.027 | 2012.1 | 0
CC71578 | 2012 | 6 | 22 | 549.21 | -0.00451 | 54 | 0.071 | 0.262 | 549.543 | 0.037 | 2012.48 | 0
CC71578 | 2013 | 2 | 19 | 549.36 | -0.00475 | 29 | 0.071 | 0.157 | 549.588 | 0.03 | 2013.14 | 0 | 2013.24 | 549.621 | 0.069 | 2 | 0.031265 | 0.094245
CC71578 | 2013 | 5 | 1 | 549.39 | -0.00472 | 37 | 0.071 | 0.194 | 549.654 | 0.032 | 2013.34 | 0
CC71578 | 2016 | 4 | 6 | 549.21 | -0.00498 | 41 | 0.071 | 0.224 | 549.505 | 0.035 | 2016.27 | 0 | 2016.27 | 549.442 | 0.047 | 4 | 0.036164 | 0.08591
CC71578 | 2016 | 4 | 7 | 549.11 | -0.00498 | 42 | 0.071 | 0.229 | 549.41 | 0.036 | 2016.27 | 0
CC71578 | 2016 | 4 | 11 | 549.06 | -0.00455 | 56 | 0.071 | 0.273 | 549.404 | 0.038 | 2016.28 | 0
CC71578 | 2016 | 4 | 12 | 549.15 | -0.0048 | 44 | 0.071 | 0.23 | 549.451 | 0.035 | 2016.28 | 0
CC71578 | 2016 | 12 | 13 | [5 empty] | 549.659 | 0.021 | 2016.95 | 0 | 2016.99 | 549.62 | 0.028 | 5 | 0.021 | 0.078242
CC71578 | 2016 | 12 | 15 | [5 empty] | 549.628 | 0.021 | 2016.96 | 0
CC71578 | 2016 | 12 | 16 | [5 empty] | 549.605 | 0.021 | 2016.96 | 0
CC71578 | 2016 | 1 | 13 | [5 empty] | 549.585 | 0.021 | 2017.04 | 0
CC71578 | 2016 | 1 | 19 | [5 empty] | 549.623 | 0.021 | 2017.05 | 0
CA08231 | 2014 | 5 | 9 | 588.675 | -0.00476 | 29 | 0.081 | 0.157 | 588.913 | 0.03 | 2014.36 | 0 | 2014.72 | 588.923 | 0.015 | 3 | 0.033694 | 0.086905
CA08231 | 2014 | 11 | 19 | 588.645 | -0.00593 | 28 | 0.081 | 0.19 | 588.916 | 0.036 | 2014.89 | 0
CA08231 | 2014 | 11 | 24 | 588.687 | -0.00595 | 25 | 0.081 | 0.173 | 588.941 | 0.035 | 2014.9 | 0
CA08231 | 2014 | 11 | 25 | 588.48 | -0.00621 | 26 | 0.081 | 0.186 | 588.747 | 0.037 | 2014.9 | 0
CA08231 | 2015 | 3 | 17 | 588.45 | -0.0053 | 45 | 0.081 | 0.26 | 588.791 | 0.039 | 2015.21 | 0 | 2015.23 | 588.866 | 0.072 | 3 | 0.038625 | 0.09778
CA08231 | 2015 | 3 | 18 | 588.54 | -0.0055 | 42 | 0.081 | 0.253 | 588.874 | 0.039 | 2015.22 | 0
CA08231 | 2015 | 4 | 2 | 588.61 | -0.0054 | 41 | 0.081 | 0.243 | 588.934 | 0.038 | 2015.26 | 0
CA08231 | 2017 | 1 | 20 | [5 empty] | 588.932 | 0.029 | 2017.05 | 0 | 2017.08 | 588.959 | 0.103 | 5 | 0.029 | 0.096467
CA08231 | 2017 | 1 | 24 | [5 empty] | 589.025 | 0.029 | 2017.07 | 0
CA08231 | 2017 | 1 | 30 | [7 empty] | 2017.08 | 1
CA08231 | 2017 | 1 | 31 | [5 empty] | 588.814 | 0.029 | 2017.09 | 0
CA08231 | 2017 | 2 | 2 | [5 empty] | 589.037 | 0.029 | 2017.09 | 0
CA08231 | 2017 | 2 | 3 | [5 empty] | 588.985 | 0.029 | 2017.09 | 0
CB11054 | 2015 | 1 | 14 | 720 | -0.0077 | 30 | 0.121 | 0.262 | 720.383 | 0.046 | 2015.04 | 0 | 2015.42 | 720.359 | 0.058 | 4 | 0.04689 | 0.11193
CB11054 | 2015 | 1 | 16 | 719.933 | -0.0078 | 30 | 0.121 | 0.265 | 720.319 | 0.047 | 2015.04 | 0
CB11054 | 2015 | 1 | 20 | 719.851 | -0.0085 | 28 | 0.121 | 0.272 | 720.244 | 0.049 | 2015.05 | 0
CB11054 | 2015 | 1 | 21 | 719.935 | -0.0083 | 30 | 0.121 | 0.282 | 720.338 | 0.049 | 2015.06 | 0
CB11054 | 2016 | 2 | 26 | 720 | -0.00716 | 34 | 0.121 | 0.272 | 720.393 | 0.045 | 2016.15 | 0 | 2016.16 | 720.375 | 0.113 | 3 | 0.044706 | 0.12548
CB11054 | 2016 | 2 | 29 | 720.08 | -0.0069 | 36 | 0.121 | 0.276 | 720.477 | 0.045 | 2016.16 | 0
CB11054 | 2016 | 3 | 2 | 719.87 | -0.0071 | 33 | 0.121 | 0.263 | 720.254 | 0.044 | 2016.17 | 0
CB11054 | 2017 | 2 | 6 | [5 empty] | 720.087 | 0.048 | 2017.1 | 0 | 2017.12 | 720.189 | 0.138 | 4 | 0.04125 | 0.12628
CB11054 | 2017 | 2 | 7 | [5 empty] | 720.07 | 0.047 | 2017.1 | 0
CB11054 | 2017 | 2 | 12 | [7 empty] | 2017.12 | 1
CB11054 | 2017 | 2 | 15 | [5 empty] | 720.235 | 0.035 | 2017.12 | 0
CB11054 | 2017 | 2 | 16 | [5 empty] | 720.364 | 0.035 | 2017.13 | 0
CB11054 | 2017 | 2 | 23 | [7 empty] | 2017.15 | 1
CC71605 | 2016 | 2 | 11 | 791.17 | -0.0077 | 34 | 0.146 | 0.293 | 791.609 | 0.048 | 2016.11 | 0 | 2016.13 | 791.415 | 0.137 | 4 | 0.046997 | 0.13548
CC71605 | 2016 | 2 | 16 | 790.99 | -0.0075 | 33 | 0.146 | 0.278 | 791.414 | 0.047 | 2016.13 | 0
CC71605 | 2016 | 2 | 17 | 790.89 | -0.0076 | 34 | 0.146 | 0.289 | 791.325 | 0.048 | 2016.13 | 0
CC71605 | 2016 | 2 | 22 | 790.89 | -0.00725 | 34 | 0.146 | 0.276 | 791.312 | 0.046 | 2016.14 | 0
CC71605 | 2017 | 2 | 27 | [5 empty] | 791.778 | 0.039 | 2017.16 | 0 | 2017.17 | 791.618 | 0.157 | 6 | 0.04 | 0.13102
CC71605 | 2017 | 2 | 28 | [5 empty] | 791.585 | 0.039 | 2017.16 | 0
CC71605 | 2017 | 3 | 1 | [5 empty] | 791.334 | 0.039 | 2017.17 | 0
CC71605 | 2017 | 3 | 2 | [5 empty] | 791.709 | 0.039 | 2017.17 | 0
CC71605 | 2017 | 3 | 3 | [5 empty] | 791.595 | 0.039 | 2017.17 | 0
CC71605 | 2017 | 3 | 6 | [5 empty] | 791.705 | 0.039 | 2017.18 | 0
CC71605 | 2020 | 1 | 10 | [5 empty] | 791.577 | 0.039 | 2020.03 | 0 | 2020.04 | 791.598 | 0.095 | 3 | 0.039 | 0.12644
CC71605 | 2020 | 1 | 17 | [5 empty] | 791.701 | 0.039 | 2020.05 | 0
CC71605 | 2020 | 1 | 20 | [5 empty] | 791.515 | 0.04 | 2020.06 | 0