## Worksheet 1
Instrument name | Brief description | Water background | Drop size | Drops/expt. | References
using reported
protocol (∘C)
Flow cell microscopy technique for aerosol phase transitions | Vapors condensed onto the bottom of a sample cell on aluminum cooling block; freezing monitored via microscope | −37 | 7–33 µm diam. | ∼65 | ,
Droplet freezing technique (DFT) | Particles deposited on a glass slide in a sample cell on a cold stage, with droplets grown by water vapor; freezing monitored via microscope | −37 | 120 µm diam. |  | ,
Microfluidic apparatus | Flow-focusing nozzle continuously produces droplets in a stream of fluorocarbon across a seven-temperature-zone cold plate; freezing monitored via microscope | −37 | 80 µm diam. | >10 000
FRankfurt Ice Deposition freezinG Experiment – Tel Aviv University (FRIDGE-TAU) | Pipetted drops onto Vaseline-coated Peltier cold stage in low-pressure diffusion chamber; freezing monitored via CCD camera | −35 | 2 µL | ∼ 120 | ,
Picoliter and Nanoliter Nucleation by Immersed Particle Instrument (pico-NIPI, nano-NIPI) | Nebulized droplets encased in silicon oil on hydrophobic glass slides on aluminum cold stage; freezing monitored via microscope | −37 | 0.25–1.7 pL; 0.1–6 nL | ∼ 135 (pL); ∼ 51 (nL) | , ,
Vienna Optical Droplet Crystallization Analyzer (VODCA) | Water–oil emulsion pipetted onto a glass slide on a Peltier cold stage, all contained in an airtight cell; freezing monitored via microscope | −36 | 10–200 µm diam. | n.f.
Drop freezing apparatus for filters | Filter cutouts placed inside small tubes with water, cooled in a water bath; freezing monitored by manual inspection | −12 | 0.1 mL | 108
Microliter Nucleation by Immersed Particle Instrument (microL-NIPI) | Drops pipetted onto a hydrophobic glass slide in humidity-controlled enclosure on a cold stage; freezing monitored via camera | −26 | 1 µL | 40 | ,
North Carolina State University cold stage (NC State-CS) | Emulsion of water in squalene placed on a glass slide resting in an aluminum dish on a thermoelectric element; freezing monitored via camera | −34 to −36 | 400 pL–150 nL | 300–1500 | ,
Microfluidic device for homogeneous ice nucleation | Microfluidically produced water-in-oil emulsion on cryo-microscopy cold stage; freezing monitored via microscope (alternatively frozen with differential scanning calorimetry) | −36 to −37 | 53–96 µm diam. | > 1000
LED-based Ice Nucleation Detection Apparatus (LINDA) | Sample in tubes held in polycarbonate tray atop an LED array submersed in a water–glycerin cooling bath; freezing monitored via camera | −15 | 40–400 µL | 52
Colorado State University Ice Spectrometer (CSU-IS) | Sample aliquots pipetted into two 96-well PCR trays cooled on custom cold blocks with N2 flow; freezing monitored via camera | −25 | 50 µL | 192 | , ,
Bielefeld Ice Nucleation ARraY (BINARY) | Droplets pipetted onto a glass slide with separated compartments atop a Peltier cold stage, all enclosed in a, N2-purged chamber; freezing monitored via camera | n.f. | 0.5–5 µL | 36
Water-Activity-Controlled Immersion Freezing Experiment (WACIFE) | Droplets pipetted onto a glass plate in a humidity-controlled aerosol conditioning cell, then sealed from ambient air and cooled on a cold stage; freezing monitored via microscope | −37 | 60–129 µm diam. | 30–50
National Institute of Polar Research Cryogenic Refrigerator Applied to Freezing Test (NIPR-CRAFT) | Drops pipetted onto a Vaseline-coated aluminum plate cooled on a cryogenic refrigerator stage; freezing monitored via camera | −33 | 5 µL | 49
Karlsruhe Institute of Technology Cold Stage (KIT-CS) | Droplets printed on a silicon substrate by piezo-driven drop-on-demand generator, and drops then covered in silicone oil and placed on cold stage; freezing monitored via CCD camera | −36 | 215 ± 70 pL | ≥ 1500
Microplate partially submersed in cooling liquid | Droplets contained in 96-well microplates partially submersed in a cooling water–alcohol bath; freezing monitored via infrared camera | −17.3 | 150 µL | 96–768
Carnegie Mellon University Cold Stage (CMU-CS) | Droplets of water in oil on a substrate in aluminum chamber cooled with a thermoelectric element; freezing monitored via microscope | −27 to −28 | 1 or 0.1 µL | 30–40 | ,
Microfluidic device and cold stage | Microfluidically generated drops in oil on a glass slide on a cryostage; freezing monitored via microscope camera | −37 | 35 µm diam. | 200
Automated Ice Spectrometer (AIS) | Drops in two 96-well PCR trays fitted into aluminum blocks fixed in a liquid cooling bath, all enclosed in an acrylic box; freezing monitored via camera | −25 to −27 | 50 µL | 192
National Oceanic and Atmospheric Administration Drop Freezing Cold Plate (NOAA-DFCP) | Drops pipetted onto Vaseline-coated copper disk placed on a thermoelectric cold plate and covered in a plastic dome; freezing monitored optically | −30 | 2.5 µL | 100
Peking University Ice Nucleation Array (PKU-INA) | Drops pipetted into compartments on a glass slide atop a cold stage in a N2-purged box; freezing monitored via CCD camera | −26 | 1 µL | 90
WeIzmann Supercooled Droplets Observation on Microarray (WISDOM) | Microfluidically produced droplet array on a PDMS surface placed on a cryostage purged with N2; freezing monitored via microscope camera | −36 | 40 or 100 µm diam. | 550 (40 µm) or 120 (100 µm)
Twin-plate Ice Nucleation Assay (TINA) | Droplets contained in four multiwell plates (2×96 and 2×384) placed on two custom aluminum cooling blocks; freezing monitored via infrared camera | −25(96-well), −28 (384-well) | 3 µL | 960
Freezing on a chip | Drops loaded on a silicon plate with etched cavities and set on thermoelectric cooler in an N2-flushed cell; freezing monitored via camera | −37.5 | 4–300 pL | 25
InfraRed Nucleation by Immersed Particles Instrument (IR-NIPI) | Drops pipetted into 96-well plate on a cold stage enclosed in a chamber; freezing monitored via infrared camera | −22 | 50 µL | 96
Ice Nucleation Droplet Array (INDA) | Samples placed in wells of a 96-well PCR tray cooled in a cooling bath; freezing monitored via CCD camera | −25 | 50 µL | 96
Leipzig Ice Nucleation Array (LINA) | Droplets pipetted into compartments on a glass slide and cooled on a Peltier element; freezing monitored via CCD camera | −30 | 1 µL | 90
Microfluidic droplet freezing assay | Microfluidically produced droplets in oil collected in microwells on glass slides placed on a Peltier cold stage in an airtight chamber; freezing monitored via microscope camera | −35 | 83–99 µm diam. | 250–500
Drop freeze assay experiment directly on exposed filters | Droplets pipetted onto filters placed on a glass slide and cold stage in an N2-purged chamber; freezing monitored via camera | −30 | 1 µL | ≤130
West Texas Cryogenic Refrigerator Applied to Freezing Test (WT-CRAFT) | Drops pipetted onto Vaseline-coated aluminum plate and cooled on a cryogenic refrigerator stage; freezing monitored via camera | −26 | 3 µL | 49
DRoplet Ice Nuclei Counter Zurich (DRINCZ) | Droplets pipetted into 96-well PCR tray submersed in ethanol cooling bath; freezing monitored via camera | −22.5 | 50 µL | 96
Cold stage to detect the most active INP in single crystals | Ice crystals placed with ultrapure water on a copper cold plate, melted, and refrozen; freezing monitored via manual inspection | −25 | 3 µL | 4
“Store and create” microfluidic device | Microfluidically generated droplets in oil in microwells of a PDMS chip placed on a cold plate sealed with an acrylic lid; freezing monitored via microscope camera | −34 | 6 nL | ≤ 720
Pyroelectric thermal sensor for ice nucleation | Drops pipetted onto Vaseline-coated pyroelectric polymer atop a cooling block; freezing monitored via pyroelectric thermal sensor | −27 | 1 µL | 30
University of Toronto Drop Freezing Technique (UT-DFT) | Drops pipetted into multiwell PCR trays cooled in an ethylene glycol water bath; freezing monitored via camera | −23 | 50 µL | 48
Ice Nucleation SpEctrometer of the Karlsruhe Institute of Technology (INSEKT) | Drops pipetted into two 96-well PCR trays cooled in custom cooling blocks; freezing monitored via camera | −20 | 50 µL | 192
Lab-On-a-Chip Nucleation by Immersed Particle Instrument (LOC-NIPI) | Water-in-oil droplets mirofluidically generated in continuous flow and passed over a series of Peltier cold plates in an N2-purged container; freezing monitored via microscope camera | −36 | 80–100 µm diam. | > 1000
Freezing Ice Nuclei Counter (FINC) | Drops pipetted into three 96-well Piko PCR trays submersed in ethanol cooling bath; freezing monitored via camera | −25 | 5–60 µL | 288 | this work