## Worksheet 1
| MC-LWCC | CSU PILS-LWCC | NOAA PILS-LWCC
Research institution | University of New Hampshire Georgia Institute of Technology | Colorado State University | NOAA Chemical SciencesLaboratory
Field campaign | FIREX-AQ 2019 | WE-CAN 2018 | FIREX-AQ 2019
Aircraft | NASA DC-8 | NSF C-130 | NOAA Twin Otter
Aircraft altitude | 200–1000 hPa | 430–1000 hPa | 630–1000 hPa
Inlet pressure control | None | None | 620 hPa
Aerosol collection | Mist chamber (Scheuer et al., 2003) | Particle-into-liquid sampler | Particle-into-liquid sampler (PILS-4001, Brechtel)
Liquid transfer | Two syringe pumps | Syringe pumps and a peristaltic pump | Peristaltic pump
Solvent | Water | Water | Water
Aerosol collection efficiencyand system dilution | Standard addition of trifluoroacetic acid to determineevaporative loss in mist chamber | Dilution ratio due to condensation was obtained from previous measurements | Aerosol collection efficiency of 0.8 for Dp
Removal of gas-phase VOCs | NA | Parallel plate carbon filter denuder | Parallel plate carbon filter denuder (DN-100, Sunset Laboratory)
Light source | Deuterium and halogen lamps (DH-mini, Ocean Optics) | Deuterium and halogen lamps (DH-mini, Ocean Optics) | Deuterium and halogen lamps (DH-mini, Ocean Optics)
Liquid waveguide | 2.5 m (LWCC-3250, World Precision Inst.) | 2.5 m (LWCC-3250, World Precision Inst.) | 2.5 m (LWCC-3250, World Precision Inst.)
Spectrometer | Ocean Optics FLAME-T-UV-VIS | Ocean Optics FLAME-T-UV-VIS | Ocean Optics QE Pro
Spectral range | 300–700 nm | 300–700 nm | 309–682 nm
Spectral resolution | 1.4 nm | 1.4 nm | 3.3 nm
Zero measurement | NA | Filtered air measured for10 min manually twice perflight | Filtered air measured for 6 min every 1.5 h
Detection limit (3σ) at365 nm | 1.53 Mm−1 | 0.89 Mm−1 | 0.03 Mm−1
Uncertainty | 28 % | 12 % | 11 %