## Worksheet 1
| MiniDOAS using AGM | HT using EC
Field performance 1,2 | 100 % uptime, 65 % flux, 35 % calibration (% calibration time can be reduced) Hands-off for 7 weeks during campaign Typical hardware maintenance: annual lamp and mirror replacement Ready for long-term monitoring deployment | 79 % uptime, no measurements during rain or dew Needs regular operator interventions (mirror cleaning) Substantial mirror degradation, replacement needed after 5 weeks4 Not yet ready for long-term monitoring deployment
Flexibility inapplication 1,3 | Needs housing and mains power Needs stable base, needs 10–20 m path Application above tall vegetation not possible yet | Can be operated on battery or solar power Relatively small, mast-mountable Remote sites and forests possible
Performance onconcentrationmeasurements | Operational in Dutch air quality monitoring network Traceable to certified standard within 3 % | Substantial deviations, correlated to ambienttemperature
Performance onflux measurements | No gold standard available, instruments compare well, with exceptions: nighttime deposition fluxes similar, miniDOAS observes higher emissions during daytime Both need ∼ 100 m homogeneous, obstacle-free upwind terrain, both need further work on analysis algorithms
| Footprint needs further study | Flux may or may not be affected by error inconcentration measurement