## ReadMe
H3O+ Ionization Product Ion Distribution (PID) Library
Last updated: 01.02.2025
Version Log
V7 | Includes data from seven laboratories, all Vocus instruments
User Community Input
Users wishing to submit data to this library can email Michael Link (michael.f.link@nist.gov) and a link to submit a data file will be provided in a follow-up email.
You will be asked to submit a chromatogram of a sample measurement from a standard calibration cylinder in the form of an unprocessed .h5 file.
Additionally, please provide the informtion corresponding to columns B through P in the "2_Lab_ID" spreadsheet.
Data submitted by users will remain anonymized in updated versions of the H3O+ PID library and follow the "Lab#" identifier formatting.
What's In This Library?
As outlined in our manuscript "Product Ion Distributions using H3O+ PTR-ToF-MS: Mechanisms, Transmission Effects, and Instrument-to-Instrument Variability" unintended product ions can form when measuring VOCs using H3O+ ionization PTR-MS.
This library contains measurements of PIDs from VOCs as determined from GC-PTR-ToF-MS measurements.
When referencing this library please reference our manuscript "Product Ion Distributions using H3O+ PTR-ToF-MS: Mechanisms, Transmission Effects, and Instrument-to-Instrument Variability", EGUsphere, https://doi.org/10.5194/egusphere-2024-3132, 2024."
Disclaimer
Certain equipment, instruments, or materials, commercial or non-commercial, are identified in this paper in to specify the experimental procedure adequately.
Such identification is not intended to imply recommendation or endorsement of any product or service by NIST, nor is it intended to imply that the materials or equipment identified are necessarily the best available for the purpose.
How to Use This Library
There are three tabs for user interaction in this spreadsheet.
1_PID_Main: The PID_Main tab contains rows of VOCs.
That spreadsheet is seperated columns differentiated by color and are described below.
2_Lab_ID: This sheet contains instrument configurations and data collection dates corresponding to the "Lab" column label in the 1_PID_Main sheet.
3_EN_Calculator: A calculator for E/N. This calculator does not incorperate additional energy from the RF-only quadrupole on the IMR of Vocus instruments in the E/N calculation.
1_PID_Main blue shading: These columns are laboratory identifier, VOC property, and measurement details
Column label | Description
Lab | Which laboratory the measurement came from. The details of instrument configurations for the laboratory ID are shown in the "Lab_ID" tab.
Species | chemical name
funtionality 1 (arbitrary assignment) | functional group 1
funtionality 2 (arbitrary assignment) | functional group 2
source | chemical source; cylinder = multi-component VOC standard cylinder, liquid = evaporated liquid source, ambient = ambient air sample
1_PID_Main green shading: These columns are elemental formula and m/q correpsonding to measured product ions elemental formula are seperated from m/q with ";" (e.g., H3O+;19.01)
Ion m/q are in units of Thompson (1 Th approximately equal to 1.036 x 10^-8 kg C^-1)
Product ion definitions are shown in the label below.
PID label | Product ion mechanism
MH+ | H+ adduct
[MH·H2O]+ | single water cluster
[MH·(H2O)2]+ | double water cluster
M+ | charge transfer
[M-H]+ | hydride transfer
[MH-H2O]+ | dehydration
Fn | fragment (n=1,2,3,etc.)
other | unknown mechanism
1_PID_Main pink shading: These columns are the product ion contributions correpsonding to measured product ions
The product ion elemental formula and m/q that appears in the green columns for a VOC will have a number between 0 and 1 at the corresponding product ion in these columns.
The sum of each product ion contribution will add up to 1 and thus the entries in these columns represent the fraction of signal attributable to a product ion generated from a VOC.
Recommended reporting uncertainties are shown in the table below. Details on the derivation of these values are in the manuscript.
Product Ion Fractional Contribution to PID | Recommended Reporting Uncertainty
> 0.30 | 15%
0.16 to 0.30 | 20%
0.04 to 0.15 | 30%
< 0.04 | 100%

## 1_PID_Main
Lab | Species | funtionality 1 | funtionality 2 | source | MH+ | [MH·H2O]+ | [MH·(H2O)2]+ | M+ | [M-H]+ | [MH-H2O]+ | F1 | F2 | F3 | F4 | F5 | other | MH+ | [MH·H2O]+ | [MH·(H2O)2]+ | M+ | [M-H]+ | [MH-H2O]+ | F1 | F2 | F3 | F4 | F5 | other
Lab1b | acetaldehyde | aldehyde |  | Liquid | C2H5O+;45.033 | C2H7O2+;63.044 | [10 empty] | 0.75 | 0.25
Lab1b | methanol | alcohol |  | Liquid | CH5O+;33.033 | CH7O2+;51.044 | CH9O3+;69.055 | [9 empty] | 0.01 | 0.91 | 0.08
Lab1b | propanal | aldehyde |  | Liquid | C3H7O+;59.049 | C3H9O2+;77.06 |  | C3H6O+;58.041 | C3H5O+;57.033 | [7 empty] | 0.31 | 0.45 |  | 0.03 | 0.21
Lab1b | furfural | aromatic | aldehyde | Cylinder | C5H5O2+;97.028 | C5H7O3+;115.04 |  | C5H4O2+;96.021 | C5H3O2+;95.013 |  | C4H5O+;69.033 | [5 empty] | 0.73 | 0.18 |  | 0.03 | 0.05 |  | 0.01
Lab1b | ethanol | alcohol |  | Liquid | C2H7O+;47.049 | C2H9O2+;65.06 | C2H11O3+;83.07 |  | C2H5O+;45.033 | [7 empty] | 0.48 | 0.46 | 0.02 |  | 0.04
Lab1b | methyl acetate | ester |  | Liquid | C3H7O2+;75.044 | C3H9O3+;93.055 |  |  |  |  | C2H5O2+;61.028 | C2H3O+;43.018 |  |  |  |  | 0.9 | 0.06 |  |  |  |  | 0.02 | 0.02
Lab1b | propanol | alcohol |  | Liquid | C3H9O+;61.065 | C3H11O2+;79.075 | C3H13O3+;97.086 |  | C3H7O+;59.049 | C3H7+;43.054 | C3H5+;41.039 |  |  |  |  | C3H9O2+;77.06 | 0.06 | 0.38 | 0.04 |  | 0.28 | 0.13 | 0.05 |  |  |  |  | 0.06
Lab1b | ethyl acetate | ester |  | Liquid | C4H9O2+;89.06 | C4H11O3+;107.07 |  |  |  |  | C2H3O+;43.018 | C2H5O2+;61.028 |  |  |  | C2H7O3+;79.039 | 0.52 | 0.07 |  |  |  |  | 0.05 | 0.31 |  |  |  | 0.05
Lab1b | butanol | alcohol |  | Liquid |  | C4H13O2+;93.091 |  |  |  | C4H9+;57.07 | [7 empty] | 0.05 |  |  |  | 0.95
Lab1b | 2-butanone | ketone |  | Liquid | C4H9O+;73.065 | C4H11O2+;91.075 | [10 empty] | 0.88 | 0.12
Lab1b | pentanal | aldehyde |  | Liquid | C5H11O+;87.08 | C5H13O2+;105.09 |  |  |  | C5H9+;69.07 | C3H5O+;57.033 | C3H7O+;59.049 |  |  |  |  | 0.1 | 0.22 |  |  |  | 0.61 | 0.04 | 0.03
Lab1b | 2,3-butadione | ketone |  | Liquid | C4H7O2+;87.044 | C4H9O3+;105.05 | C4H11O4+;123.07 |  |  |  | C3H7O+;59.049 | [5 empty] | 0.73 | 0.12 | 0.03 |  |  |  | 0.12
Lab1b | octanal | aldehyde |  | Liquid | C8H17O+;129.13 | C8H19O2+;147.14 |  |  | C8H15O+;127.11 | C8H15+;111.12 | C5H9+;69.07 | C3H5+;41.039 |  |  |  |  | 0.34 | 0.17 |  |  | 0.02 | 0.25 | 0.19 | 0.03
Lab1b | butanal | aldehyde |  | Liquid | C4H9O+;73.065 | C4H11O2+;91.075 |  | C4H8O+;72.057 | C4H7O+;71.049 | C4H7+;55.054 | C4H5+;53.039 | C3H5O+;57.033 |  |  |  |  | 0.23 | 0.06 |  | 0.03 | 0.01 | 0.6 | 0.04 | 0.03
Lab1b | linalool | alkene | alcohol | Liquid | C10H19O+;155.14 |  |  |  |  | C10H17+;137.13 | C6H9+;81.07 | C3H7O+;59.049 | C7H11+;95.086 |  |  |  | 0.03 |  |  |  |  | 0.48 | 0.42 | 0.04 | 0.03
Lab1b | 6-methyl-5-heptene-2-one | ketone | alkene | Liquid | C8H15O+;127.11 |  |  |  |  | C8H13+;109.1 | C8H12+;108.09 | [5 empty] | 0.52 |  |  |  |  | 0.44 | 0.04
Lab1b | formic acid | acid |  | Liquid | CH3O2+;47.013 | CH5O3+;65.023 | [10 empty] | 0.47 | 0.53
Lab1b | acetic acid | acid |  | Liquid | C2H5O2+;61.028 | C2H7O3+;79.039 |  |  |  | C2H3O+;43.018 | [6 empty] | 0.4 | 0.13 |  |  |  | 0.47
Lab1b | propanoic acid | acid |  | Liquid | C3H7O2+;75.044 | C3H9O3+;93.055 |  | C3H6O2+;74.036 |  | C3H5O+;57.033 | [6 empty] | 0.57 | 0.12 |  | 0.04 |  | 0.27
Lab1b | butanoic acid | acid |  | Liquid | C4H9O2+;89.06 | C4H11O3+;107.07 |  |  |  | C4H7O+;71.049 | C3H5+;41.039 | C3H7+;43.054 | C2H4O2+;60.021 |  |  |  | 0.45 | 0.1 |  |  |  | 0.19 | 0.06 | 0.17 | 0.03
Lab1b | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.74 | 0.26
Lab1b | pentanoic acid | acid |  | Liquid | C5H11O2+;103.08 | C5H13O3+;121.09 | C5H15O4+;139.1 |  |  | C5H9O+;85.065 | C4H9+;57.07 | C2H4O2+;60.021 | C3H5+;41.039 | C3H3+;39.023 |  |  | 0.12 | 0.34 | 0.02 |  |  | 0.17 | 0.32 | 0.01 | 0.01 | 0.01
Lab1b | unidentified monoterpene acetate ester | ester | alkene | Ambient | C12H21O2+;197.15 | [5 empty] | C10H17+;137.13 | C6H9+;81.07 | C2H5O+;45.033 | C3H7O+;59.049 | C2H7O2+;63.044 |  | 0.01 | [5 empty] | 0.14 | 0.14 | 0.48 | 0.06 | 0.17
Lab1a | heptanal | aldehyde |  | Cylinder | C7H15O+;115.11 | C7H17O2+;133.12 |  |  | C7H13O+;113.1 | C7H13+;97.101 | C5H9+;69.07 | C3H5O+;57.033 | C4H7+;55.054 |  |  |  | 0.12 | 0.18 |  |  | 0.02 | 0.51 | 0.02 | 0.01 | 0.14
Lab1a | hexanal | aldehyde |  | Cylinder | C6H13O+;101.1 | C6H15O2+;119.11 |  |  | C6H11O+;99.08 | C6H11+;83.086 | C4H7+;55.054 | [5 empty] | 0.07 | 0.19 |  |  | 0.01 | 0.57 | 0.16
Lab1a | cyclohexene | alkene |  | Cylinder | C6H11+;83.086 |  |  | C6H10+;82.078 | C6H9+;81.07 |  | C4H5+;53.039 | C5H7+;67.054 |  |  |  |  | 0.78 |  |  | 0.03 | 0.03 |  | 0.01 | 0.15
Lab1a | hydroxyacetone | ketone | alcohol | Cylinder | C3H7O2+;75.044 | C3H9O3+;93.055 |  |  |  | C3H5O+;57.033 | [6 empty] | 0.9 | 0.02 |  |  |  | 0.08
Lab1a | tetrahydrofuran | heterocycle |  | Cylinder | C4H9O+;73.065 | C4H11O2+;91.075 |  |  | C4H7O+;71.049 | C4H7+;55.054 | C4H5+;53.039 | [5 empty] | 0.37 | 0.33 |  |  | 0.04 | 0.25 | 0.01
Lab1a | methacrolein | aldehyde | alkene | Cylinder | C4H7O+;71.049 | C4H9O2+;89.06 | [10 empty] | 0.61 | 0.39
Lab1a | nitromethane | nitro |  | Cylinder | CH4NO2+;62.024 | CH6NO3+;80.034 | [10 empty] | 0.93 | 0.07
Lab1a | dimethylsulfide | sulfide |  | Cylinder | C2H7S+;63.026 | C2H6S+;62.018 | [10 empty] | 0.96 | 0.04
Lab1a | acrylonitrile | nitrile | alkene | Cylinder | C3H4N+;54.034 | C3H6NO+;72.044 | [10 empty] | 0.9 | 0.1
Lab1a | acetone | ketone |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 | [10 empty] | 0.9 | 0.1
Lab1a | citral | aldehyde | alkene | Cylinder | [5 empty] | C10H15+;135.12 | C7H9+;93.07 | C8H11+;107.09 | C7H7+;91.054 | C9H11+;119.09 | [7 empty] | 0.38 | 0.39 | 0.06 | 0.08 | 0.09
Lab1a | phenol | aromatic | alcohol | Cylinder | C6H7O+;95.049 |  |  | C6H6O+;94.041 | [8 empty] | 0.89 |  |  | 0.11
Lab1a | styrene | aromatic | alkene | Cylinder | C8H9+;105.07 |  |  | C8H8+;104.06 | [8 empty] | 0.87 |  |  | 0.13
Lab1a | o-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 |  |  | C7H7+;91.054 | [5 empty] | 0.81 |  |  | 0.11 |  |  | 0.08
Lab1a | chlorobenzene | aromatic | halogen | Cylinder | C6ClH6+;113.02 |  |  | C6ClH5+;112.01 | [7 empty] | C6H7O+;95.049 | 0.68 |  |  | 0.23 | [7 empty] | 0.09
Lab1a | benzene | aromatic |  | Cylinder | C6H7+;79.054 |  |  | C6H6+;78.046 | [7 empty] | C6H7O+;95.049 | 0.74 |  |  | 0.2 | [7 empty] | 0.06
Lab1a | D5 | siloxane | cyclic | Cylinder | C10H31O5Si5+;371.1 | [5 empty] | C9H27O5Si5+;355.07 | [5 empty] | 0.96 | [5 empty] | 0.04
Lab1a | nonanal | aldehyde |  | Cylinder | C9H19O+;143.14 | C9H21O2+;161.15 |  |  |  | C9H17+;125.13 | C6H11+;83.086 | C5H9+;69.07 |  |  |  |  | 0.3 | 0.3 |  |  |  | 0.1 | 0.11 | 0.19
Lab1a | pyrolle | heterocycle | alkene | Cylinder | C4H6N+;68.049 | [11 empty] | 1
Lab1a | 2-pentanone | ketone |  | Cylinder | C5H11O+;87.08 | C5H13O2+;105.09 |  |  |  | C5H9+;69.07 | C2H5O+;45.033 | [5 empty] | 0.64 | 0.32 |  |  |  | 0.02 | 0.02
Lab1a | 2-octanone | ketone |  | Cylinder | C8H17O+;129.13 | C8H19O2+;147.14 |  |  |  |  | C9H13+;121.1 | C5H9+;69.07 |  |  |  |  | 0.69 | 0.25 |  |  |  |  | 0.04 | 0.02
Lab1a | butanal | aldehyde |  | Cylinder | C4H9O+;73.065 | C4H11O2+;91.075 |  | C4H8O+;72.057 | C4H7O+;71.049 | C4H7+;55.054 | C4H5+;53.039 | C3H5O+;57.033 |  |  |  |  | 0.1 | 0.32 |  | 0.01 | 0.01 | 0.52 | 0.02 | 0.02
Lab1a | isoprene | alkene |  | Cylinder | C5H9+;69.07 |  |  | C5H8+;68.062 | C5H7+;67.054 | [7 empty] | 0.84 |  |  | 0.06 | 0.1
Lab1a | 1,1-dichloroethane | halogen |  | Cylinder | [6 empty] | C2H3Cl+;61.992 | [11 empty] | 1
Lab1a | 1,3-hexachlorobutadiene | halogen | alkene | Cylinder | C4HCl6+;258.82 |  |  | C4Cl6+;257.81 | [8 empty] | 0.43 |  |  | 0.57
Lab1a | 1,2-dichlorobenzene | aromatic | halogen | Cylinder | C6Cl2H5+;146.98 |  |  | C6H4Cl2+;145.97 | [8 empty] | 0.65 |  |  | 0.35
Lab1a | 1,3-dichlorobenzene | aromatic | halogen | Cylinder | C6Cl2H5+;146.98 |  |  | C6H4Cl2+;145.97 | [7 empty] | C6ClH6O+;129.01 | 0.73 |  |  | 0.24 | [7 empty] | 0.03
Lab1a | camphene | cyclic | alkene | Cylinder | C10H17+;137.13 |  |  | C10H16+;136.12 |  |  | C6H9+;81.07 | C9H13+;121.1 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 |  | 0.51 |  |  | 0.02 |  |  | 0.35 | 0.04 | 0.04 | 0.03 | 0.01
Lab1a | b-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 |  |  | C10H16+;136.12 |  |  | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 |  |  | 0.47 |  |  | 0.02 |  |  | 0.4 | 0.06 | 0.03 | 0.02
Lab1a | α-terpinene | cyclic | alkene | Cylinder | C10H17+;137.13 |  |  | C10H16+;136.12 |  |  | C6H9+;81.07 | C9H13+;121.1 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 |  | 0.48 |  |  | 0.07 |  |  | 0.34 | 0.05 | 0.03 | 0.02 | 0.01
Lab1a | p-cymene | cyclic | alkene | Cylinder | C10H15+;135.12 |  |  | C10H14+;134.11 |  |  | C7H9+;93.07 | C9H11+;119.09 | C7H7+;91.054 |  |  |  | 0.21 |  |  | 0.04 |  |  | 0.55 | 0.14 | 0.06
Lab1a | α-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C7H7+;91.054 | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 |  | 0.47 | [5 empty] | 0.03 | 0.38 | 0.07 | 0.03 | 0.02
Lab1a | delta-3-carene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 | C7H8+;92.062 |  | 0.54 | [5 empty] | 0.33 | 0.05 | 0.03 | 0.03 | 0.02
Lab1a | limonene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H11+;95.086 | C6H7+;79.054 |  |  |  | 0.45 | [5 empty] | 0.44 | 0.07 | 0.04
Lab1b | 1,8-cineole | cyclic | alcohol | Cylinder | C10H19O+;155.14 |  |  |  |  | C10H17+;137.13 | C6H9+;81.07 | C7H11+;95.086 |  |  |  |  | 0.06 |  |  |  |  | 0.53 | 0.39 | 0.02
Lab1a | trichloroethylene | halogen | alkene | Cylinder |  |  |  | C2HCl3+;129.91 | [11 empty] | 1
Lab1a | carbontetrachloride | halogen |  | Cylinder | [6 empty] | CCl3+;116.91 | [11 empty] | 1
Lab1a | toluene | aromatic |  | Cylinder | C7H9+;93.07 |  |  | C7H8+;92.062 | C7H7+;91.054 | [6 empty] | C6H7O+;95.049 | 0.73 |  |  | 0.1 | 0.13 | [6 empty] | 0.04
Lab1a | tetrachloroethylene | halogen | alkene | Cylinder |  |  |  | C2Cl4+;163.87 | [11 empty] | 1
Lab1a | 1,4-dichlorobenzene | aromatic | halogen | Cylinder | C6Cl2H5+;146.98 |  |  | C6H4Cl2+;145.97 | [6 empty] | C6ClH6O+;129.01 |  | 0.65 |  |  | 0.34 | [7 empty] | 0.01
Lab1a | trichlorofluoromethane | halogen |  | Cylinder | [6 empty] | CCl3+;116.91 | CFCl2+;100.94 | [10 empty] | 0.61 | 0.39
Lab1a | vinyl chloride | halogen | alkene | Cylinder |  |  |  | C2H3Cl+;61.992 | [11 empty] | 1
Lab1a | ethyl benzene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 |  |  | C7H7+;91.054 | C6H7+;79.054 |  |  |  |  | 0.47 |  |  | 0.07 |  |  | 0.1 | 0.36
Lab1a | methylcyclopentanone | cyclic | ketone | Liquid | C6H9O+;97.065 | C6H11O2+;115.08 | [10 empty] | 0.91 | 0.09
Lab1a | cyclopentanone | cyclic | ketone | Liquid | C5H7O+;83.049 | C5H9O2+;101.06 |  | C5H6O+;82.041 | [8 empty] | 0.88 | 0.09 |  | 0.03
Lab1a | propylene glycol | alcohol |  | Liquid | C3H9O2+;77.06 | C3H11O3+;95.07 |  |  | C3H7O2+;75.044 | C3H7O+;59.049 | [6 empty] | 0.32 | 0.04 |  |  | 0.1 | 0.54
Lab1a | 2-hexanone | ketone |  | Liquid | C6H13O+;101.1 | C6H15O2+;119.11 |  |  |  | C6H11+;83.086 | [6 empty] | 0.52 | 0.4 |  |  |  | 0.08
Lab1a | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.4 | 0.6
Lab1a | 2-butanone | ketone |  | Cylinder | C4H9O+;73.065 | C4H11O2+;91.075 |  |  |  | C4H7+;55.054 | [6 empty] | 0.81 | 0.15 |  |  |  | 0.04
Lab2a | ethyl acetate | ester |  | Cylinder | C4H9O2+;89.06 | C4H11O3+;107.07 |  |  |  |  | C2H3O+;43.018 | C2H5O2+;61.028 |  |  |  | C2H7O3+;79.039 | 0.43 | 0.11 |  |  |  |  | 0.38 | 0.04 |  |  |  | 0.04
Lab2a | vinyl acetate | ester | alkene | Cylinder | C4H7O2+;87.044 | C4H9O3+;105.05 |  |  |  |  | C2H3O+;43.018 | C2H5O2+;61.028 |  |  |  | C2H7O3+;79.039 | 0.26 | 0.05 |  |  |  |  | 0.19 | 0.41 |  |  |  | 0.09
Lab2a | acetone | ketone |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 | [10 empty] | 0.95 | 0.05
Lab2a | 1,2,4-trichlorobenzene | aromatic | halogen | Cylinder | C6H4Cl3+;180.94 |  |  | C6H3Cl3+;179.93 | [8 empty] | 0.4 |  |  | 0.6
Lab2a | 1,2,4-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 | C9H11+;119.09 |  | C8H9+;105.07 | [5 empty] | 0.89 |  |  | 0.06 | 0.01 |  | 0.04
Lab2a | 1,2-dichlorobenzene | aromatic | halogen | Cylinder | C6H5Cl2+;146.98 |  |  | C6H4Cl2+;145.97 | [8 empty] | 0.6 |  |  | 0.4
Lab2a | 1,3,5-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 | C9H11+;119.09 |  | C8H9+;105.07 | [5 empty] | 0.91 |  |  | 0.04 | 0.02 |  | 0.03
Lab2a | 1,3-dichlorobenzene | aromatic | halogen | Cylinder | C6H5Cl2+;146.98 |  |  | C6H4Cl2+;145.97 | [8 empty] | 0.79 |  |  | 0.21
Lab2a | 1,4-dichlorobenzene | aromatic | halogen | Cylinder | C6H5Cl2+;146.98 |  |  | C6H4Cl2+;145.97 | [8 empty] | 0.65 |  |  | 0.35
Lab2a | chlorobenzene | aromatic | halogen | Cylinder | C6ClH6+;113.02 |  |  | C6ClH5+;112.01 | [7 empty] | C6H7O+;95.049 | 0.71 |  |  | 0.21 | [7 empty] | 0.08
Lab2a | ethyl benzene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 |  |  | C7H7+;91.054 | C6H7+;79.054 |  |  |  |  | 0.57 |  |  | 0.04 |  |  | 0.08 | 0.31
Lab2a | benzene | aromatic |  | Cylinder | C6H7+;79.054 |  |  | C6H6+;78.046 | [7 empty] | C6H7O+;95.049 | 0.73 |  |  | 0.16 | [7 empty] | 0.11
Lab2a | benzyl chloride | aromatic | halogen | Cylinder | C7H7+;91.054 | [11 empty] | 1
Lab2a | 1,3-hexachlorobutadiene | alkene | halogen | Cylinder | C4HCl6+;258.82 |  |  | C4Cl6+;257.81 | [8 empty] | 0.61 |  |  | 0.39
Lab2a | 2-butanone | ketone |  | Cylinder | C4H9O+;73.065 | C4H11O2+;91.075 | [10 empty] | 0.91 | 0.09
Lab2a | carbon disulfide | sulfide |  | Cylinder | CS2+;75.944 | [11 empty] | 1
Lab2a | 1,4-dioxane | heterocycle |  | Cylinder | C4H9O2+;89.06 | C4H11O3+;107.07 |  |  | C4H7O2+;87.044 |  | C2H5O+;45.033 | [5 empty] | 0.43 | 0.47 |  |  | 0.02 |  | 0.08
Lab2a | 1,1,1-trichloroethane | halogen |  | Cylinder | C2H3Cl2+;96.961 | [11 empty] | 1
Lab2a | 1,1,2,2-tetrachloroethane | halogen |  | Cylinder | [6 empty] | C2HCl3+;129.91 | CHCl2+;82.945 | [10 empty] | 0.1 | 0.9
Lab2a | 1,2-dibromomethane | halogen |  | Cylinder | [6 empty] | C2H4Br+;106.95 | [11 empty] | 1
Lab2a | 1,2-dichloroethane | halogen |  | Cylinder | [6 empty] | C2H3Cl+;61.992 | [11 empty] | 1
Lab2a | ethanol | alcohol |  | Cylinder | C2H7O+;47.049 | C2H9O2+;65.06 | C2H11O3+;83.07 | [9 empty] | 0.54 | 0.44 | 0.02
Lab2a | 1,1-dichloroethylene | halogen | alkene | Cylinder | C2H3Cl2+;96.961 |  |  | C2H2Cl2+;95.953 | [8 empty] | 0.83 |  |  | 0.17
Lab2a | cis-1,1-dichloroethylene | halogen | alkene | Cylinder | C2H3Cl2+;96.961 |  |  | C2H2Cl2+;95.953 | [8 empty] | 0.04 |  |  | 0.96
Lab2a | trans-1,2-dichloroethylene | halogen | alkene | Cylinder |  |  |  | C2H2Cl2+;95.953 | [11 empty] | 1
Lab2a | tetrachloroethylene | halogen | alkene | Cylinder |  |  |  | C2Cl4+;163.87 | [11 empty] | 1
Lab2a | trichloroethylene | halogen | alkene | Cylinder |  |  |  | C2HCl3+;129.91 | [11 empty] | 1
Lab2a | tetrahydrofuran | heterocycle | alkene | Cylinder | C4H9O+;73.065 | C4H11O2+;91.075 |  |  | C4H7O+;71.049 | [7 empty] | 0.8 | 0.17 |  |  | 0.03
Lab2a | 2-hexanone | ketone |  | Cylinder | C6H13O+;101.1 | C6H15O2+;119.11 |  |  |  | C6H11+;83.086 | [6 empty] | 0.71 | 0.25 |  |  |  | 0.04
Lab2a | bromodichloromethane | halogen |  | Cylinder | [6 empty] | CHClBr+;126.89 | CHCl2+;82.945 | [10 empty] | 0.6 | 0.4
Lab2a | dibromochloromethane | halogen |  | Cylinder | [6 empty] | CHClBr+;126.89 | CHBr2+;170.84 | [10 empty] | 0.73 | 0.27
Lab2a | carbontetrachloride | halogen |  | Cylinder | [6 empty] | CCl3+;116.91 | [11 empty] | 1
Lab2a | tribromomethane | halogen |  | Cylinder | [6 empty] | CHBr2+;170.84 | [11 empty] | 1
Lab2a | trichlorofluoromethane | halogen |  | Cylinder | [6 empty] | CCl3+;116.91 | CFCl2+;100.94 | [10 empty] | 0.73 | 0.27
Lab2a | trichloromethane | halogen |  | Cylinder | [6 empty] | CCl3+;116.91 | CHCl2+;82.945 | [10 empty] | 0.01 | 0.99
Lab2a | 2-methylpentan-4-one | ketone |  | Cylinder | C6H13O+;101.1 | C6H15O2+;119.11 |  |  |  | C6H11+;83.086 | C4H9+;57.07 | C3H7O+;59.049 |  |  |  |  | 0.66 | 0.2 |  |  |  | 0.02 | 0.08 | 0.04
Lab2a | 2-propanol | alcohol |  | Cylinder | C3H9O+;61.065 | C3H11O2+;79.075 |  |  | C3H7O+;59.049 | C3H7+;43.054 | [6 empty] | 0.11 | 0.33 |  |  | 0.35 | 0.21
Lab2a | cis-1,3-dichloropropene | halogen | alkene | Cylinder | [6 empty] | C3H4Cl+;75 | [11 empty] | 1
Lab2a | trans-1,3-dichloropropene | halogen | alkene | Cylinder | [6 empty] | C3H4Cl+;75 | [11 empty] | 1
Lab2a | styrene | aromatic | alkene | Cylinder | C8H9+;105.07 |  |  | C8H8+;104.06 | C8H7+;103.05 | [7 empty] | 0.9 |  |  | 0.08 | 0.02
Lab2a | 4-ethyltoluene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 |  |  | C7H9+;93.07 | C7H7+;91.054 | C8H9+;105.07 |  |  | C6H7O+;95.049 | 0.62 |  |  | 0.02 |  |  | 0.21 | 0.07 | 0.03 |  |  | 0.05
Lab2a | toluene | aromatic |  | Cylinder | C7H9+;93.07 |  |  | C7H8+;92.062 | C7H7+;91.054 | [6 empty] | C6H7O+;95.049 | 0.75 |  |  | 0.07 | 0.11 | [6 empty] | 0.07
Lab2a | o-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 |  |  | C7H7+;91.054 | [5 empty] | 0.86 |  |  | 0.04 |  |  | 0.1
Lab2b | ethanol | alcohol |  | Cylinder | C2H7O+;47.049 | C2H9O2+;65.06 |  |  | C2H5O+;45.033 | [7 empty] | 0.59 | 0.39 |  |  | 0.02
Lab2b | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.76 | 0.24
Lab2b | acetone | ketone |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 | [10 empty] | 0.98 | 0.02
Lab2b | acrylonitrile | nitrile | alkene | Cylinder | C3H4N+;54.034 | C3H6NO+;72.044 | [10 empty] | 0.91 | 0.09
Lab2b | isoprene | alkene |  | Cylinder | C5H9+;69.07 |  |  | C5H8+;68.062 | C5H7+;67.054 | [7 empty] | 0.92 |  |  | 0.02 | 0.06
Lab2b | 2-butanone | ketone |  | Cylinder | C4H9O+;73.065 | [11 empty] | 1
Lab2b | benzene | aromatic |  | Cylinder | C6H7+;79.054 |  |  | C6H6+;78.046 | [7 empty] | C6H7O+;95.049 | 0.92 |  |  | 0.01 | [7 empty] | 0.07
Lab2b | m-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 | C8H9+;105.07 |  | C7H7+;91.054 | [5 empty] | 0.88 |  |  | 0.03 | 0.02 |  | 0.07
Lab2b | a-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C7H7+;91.054 | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 |  | 0.5 | [5 empty] | 0.01 | 0.42 | 0.02 | 0.03 | 0.02
Lab2b | 1,2,4-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  |  | C9H11+;119.09 |  | C8H9+;105.07 | [5 empty] | 0.95 |  |  |  | 0.01 |  | 0.04
Lab2b | D3 | siloxane |  | Cylinder | C6H19O3Si3+;223.06 | C6H21O4Si3+;241.07 | [10 empty] | 0.96 | 0.04
Lab2b | D4 | siloxane |  | Cylinder | C8H25O4Si4+;297.08 | [5 empty] | C7H21O4Si4+;281.05 | [5 empty] | 0.98 | [5 empty] | 0.02
Lab2b | D5 | siloxane |  | Cylinder | C10H31O5Si5+;371.1 | [5 empty] | C9H27O5Si5+;355.07 | [5 empty] | 0.8 | [5 empty] | 0.2
Lab3a | acetone | ketone |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 |  |  |  |  | C2H3O+;43.018 | [5 empty] | 0.95 | 0.03 |  |  |  |  | 0.02
Lab3a | propanal | aldehyde |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 |  |  |  |  | C2H3O+;43.018 | [5 empty] | 0.91 | 0.04 |  |  |  |  | 0.05
Lab3a | isoprene | alkene |  | Cylinder | C5H9+;69.07 |  |  |  | C5H7+;67.054 |  | C3H5+;41.039 | C3H3+;39.023 |  |  |  |  | 0.39 |  |  |  | 0.06 |  | 0.28 | 0.27
Lab3a | methacrolein | aldehyde | alkene | Cylinder | C4H7O+;71.049 | C4H9O2+;89.06 |  |  |  |  | C3H7+;43.054 | C3H5+;41.039 | C3H3+;39.023 |  |  |  | 0.61 | 0.08 |  |  |  |  | 0.16 | 0.08 | 0.07
Lab3a | methyl vinyl ketone | ketone | alkene | Cylinder | C4H7O+;71.049 | C4H9O2+;89.06 |  |  |  |  | C2H3O+;43.018 | [5 empty] | 0.78 | 0.04 |  |  |  |  | 0.18
Lab3a | 2-methyl-3-buten-2-ol | alcohol | alkene | Cylinder | C5H11O+;87.08 | C5H13O2+;105.09 |  |  |  | C5H9+;69.07 | C3H5+;41.039 | C3H3+;39.023 |  |  |  |  | 0.1 | 0.01 |  |  |  | 0.42 | 0.26 | 0.21
Lab3a | ethyl acetate | ester |  | Cylinder | C4H9O2+;89.06 | [5 empty] | C2H3O+;43.018 | C2H5O2+;61.028 |  |  |  | C2H7O3+;79.039 | 0.11 | [5 empty] | 0.25 | 0.58 |  |  |  | 0.06
Lab3a | toluene | aromatic |  | Cylinder | C7H9+;93.07 |  |  | C7H8+;92.062 | C7H7+;91.054 | [6 empty] | C6H7O+;95.049 | 0.67 |  |  | 0.04 | 0.17 | [6 empty] | 0.12
Lab3a | a-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 |  |  | 0.32 | [5 empty] | 0.54 | 0.03 | 0.07 | 0.04
Lab3a | b-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 |  |  | C10H16+;136.12 |  |  | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 |  |  | 0.3 | [5 empty] | 0.56 | 0.02 | 0.07 | 0.05
Lab3a | limonene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H11+;95.086 | C6H7+;79.054 |  |  |  | 0.26 | [5 empty] | 0.62 | 0.11 | 0.01
Lab3a | D5 | siloxane |  | Cylinder | C10H31O5Si5+;371.1 | [5 empty] | C9H27O5Si5+;355.07 | [5 empty] | 0.83 | [5 empty] | 0.17
Lab3b | ethanol | alcohol |  | Cylinder | C2H7O+;47.049 | C2H9O2+;65.06 | [10 empty] | 0.7 | 0.3
Lab3b | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.97 | 0.03
Lab3b | acetone | ketone |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 |  |  |  |  | C2H3O+;43.018 | [5 empty] | 0.9 | 0.07 |  |  |  |  | 0.03
Lab3b | acrylonitrile | nitrile | alkene | Cylinder | C3H4N+;54.034 | C3H6NO+;72.044 | [10 empty] | 0.91 | 0.09
Lab3b | isoprene | alkene |  | Cylinder | C5H9+;69.07 |  |  | C5H8+;68.062 | C5H7+;67.054 |  | C3H5+;41.039 | C3H3+;39.023 |  |  |  |  | 0.66 |  |  | 0.01 | 0.06 |  | 0.16 | 0.11
Lab3b | 2-butanone | ketone |  | Cylinder | C4H9O+;73.065 | C4H11O2+;91.075 | [10 empty] | 0.82 | 0.18
Lab3b | methyl vinyl ketone | ketone | alkene | Cylinder | C4H7O+;71.049 | C4H9O2+;89.06 |  |  |  |  | C2H3O+;43.018 | [5 empty] | 0.87 | 0.09 |  |  |  |  | 0.04
Lab3b | benzene | aromatic |  | Cylinder | C6H7+;79.054 |  |  | C6H6+;78.046 | [7 empty] | C6H7O+;95.049 | 0.83 |  |  | 0.11 | [7 empty] | 0.06
Lab3b | m-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 | C8H9+;105.07 | [7 empty] | 0.94 |  |  | 0.04 | 0.02
Lab3b | a-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 |  |  | 0.42 | [5 empty] | 0.51 | 0.03 | 0.03 | 0.01
Lab3b | 1,2,4-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 | C9H11+;119.09 |  | C8H9+;105.07 | [5 empty] | 0.87 |  |  | 0.07 | 0.02 |  | 0.04
Lab3b | D4 | siloxane |  | Cylinder | C8H25O4Si4+;297.08 | C8H27O5Si4+;315.09 | [10 empty] | 0.93 | 0.07
Lab3b | D5 | siloxane |  | Cylinder | C10H31O5Si5+;371.1 | [5 empty] | C9H27O5Si5+;355.07 | [5 empty] | 0.79 | [5 empty] | 0.21
Lab4a | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.84 | 0.16
Lab4a | a-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 |  |  | 0.33 | [5 empty] | 0.57 | 0.02 | 0.05 | 0.03
Lab4a | m-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 | C8H9+;105.07 | [7 empty] | 0.94 |  |  | 0.02 | 0.04
Lab4a | D5 | siloxane |  | Cylinder | C10H31O5Si5+;371.1 | [5 empty] | C9H27O5Si5+;355.07 | [5 empty] | 0.89 | [5 empty] | 0.11
Lab4b | benzene | aromatic |  | Cylinder | C6H7+;79.054 |  |  | C6H6+;78.046 | [7 empty] | C6H7O+;95.049 | 0.81 |  |  | 0.08 | [7 empty] | 0.11
Lab4b | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.87 | 0.13
Lab4b | acetone | ketone |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 | [10 empty] | 0.91 | 0.09
Lab4b | m-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 | C8H9+;105.07 |  | C7H7+;91.054 | [5 empty] | 0.82 |  |  | 0.03 | 0.04 |  | 0.11
Lab4b | a-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 |  |  | 0.36 | [5 empty] | 0.53 | 0.03 | 0.05 | 0.03
Lab5 | methanol | alcohol |  | Cylinder | CH5O+;33.033 | CH7O2+;51.044 | [10 empty] | 0.78 | 0.22
Lab5 | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.97 | 0.03
Lab5 | acetaldehyde | aldehyde |  | Cylinder | C2H5O+;45.033 | C2H7O2+;63.044 | [10 empty] | 0.95 | 0.05
Lab5 | acetone | ketone |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 |  |  |  |  | C2H3O+;43.018 | [5 empty] | 0.91 | 0.06 |  |  |  |  | 0.03
Lab5 | isoprene | alkene |  | Cylinder | C5H9+;69.07 |  |  |  | C5H7+;67.054 |  | C3H5+;41.039 | [5 empty] | 0.8 |  |  |  | 0.05 |  | 0.15
Lab5 | benzene | aromatic |  | Cylinder | C6H7+;79.054 |  |  | C6H6+;78.046 | [7 empty] | C6H7O+;95.049 | 0.87 |  |  | 0.08 | [7 empty] | 0.05
Lab5 | toluene | aromatic |  | Cylinder | C7H9+;93.07 |  |  | C7H8+;92.062 | C7H7+;91.054 | [6 empty] | C6H7O+;95.049 | 0.81 |  |  | 0.05 | 0.05 | [6 empty] | 0.09
Lab5 | m-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 |  |  | C7H7+;91.054 | [5 empty] | 0.89 |  |  | 0.04 |  |  | 0.07
Lab5 | 1,2,3-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 |  |  | C8H9+;105.07 | [5 empty] | 0.92 |  |  | 0.06 |  |  | 0.02
Lab5 | a-pinene | cyclic | alkene | Ambient | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | [5 empty] | 0.49 | [5 empty] | 0.51
Lab5 | ethanol | alcohol |  | Ambient | C2H7O+;47.049 | C2H9O2+;65.06 |  |  | C2H5O+;45.033 | [7 empty] | 0.57 | 0.26 |  |  | 0.17
Lab6a | isoprene | alkene |  | Cylinder | C5H9+;69.07 |  |  | C5H8+;68.062 | C5H7+;67.054 |  | C3H5+;41.039 | C3H3+;39.023 |  |  |  |  | 0.61 |  |  | 0.08 | 0.19 |  | 0.07 | 0.05
Lab6a | acetone | ketone |  | Cylinder | C3H7O+;59.049 | [5 empty] | C2H3O+;43.018 | [5 empty] | 0.94 | [5 empty] | 0.06
Lab6a | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.99 | 0.01
Lab6a | acrylonitrile | nitrile | alkene | Cylinder | C3H4N+;54.034 | C3H6NO+;72.044 | [10 empty] | 0.99 | 0.01
Lab6a | 2-butanone | ketone |  | Cylinder | C4H9O+;73.065 |  |  |  |  | C4H7+;55.054 | C2H3O+;43.018 | [5 empty] | 0.87 |  |  |  |  | 0.09 | 0.04
Lab6a | benzene | aromatic |  | Cylinder | C6H7+;79.054 |  |  | C6H6+;78.046 | [7 empty] | C6H7O+;95.049 | 0.67 |  |  | 0.28 | [7 empty] | 0.05
Lab6a | toluene | aromatic |  | Cylinder | C7H9+;93.07 |  |  | C7H8+;92.062 | C7H7+;91.054 | [7 empty] | 0.66 |  |  | 0.22 | 0.12
Lab6a | m-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 | C8H9+;105.07 |  | C7H7+;91.054 | [5 empty] | 0.7 |  |  | 0.18 | 0.02 |  | 0.1
Lab6a | a-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H9+;93.07 | C6H7+;79.054 |  |  |  | 0.42 | [5 empty] | 0.41 | 0.15 | 0.02
Lab6a | 1,2,4-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 | C9H11+;119.09 |  | C8H9+;105.07 | [5 empty] | 0.72 |  |  | 0.2 | 0.01 |  | 0.07
Lab6a | furan | heterocycle |  | Cylinder | C4H5O+;69.033 | [11 empty] | 1
Lab6a | carbontetrachloride | halogen |  | Cylinder | [6 empty] | CCl3+;116.91 | [11 empty] | 1
Lab6a | t-Amyl Ethyl Ether | ether |  | Cylinder | [6 empty] | C5H11O+;87.08 | C5H11+;71.086 | C4H9O+;73.065 | C3H7O+;59.049 | C3H7+;43.054 | C6H13O+;101.1 | [6 empty] | 0.11 | 0.42 | 0.05 | 0.22 | 0.16 | 0.04
Lab6a | ethylbenzene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 |  |  | C7H7+;91.054 | C6H7+;79.054 |  |  |  | C6H7O+;95.049 | 0.36 |  |  | 0.07 |  |  | 0.2 | 0.35 |  |  |  | 0.02
Lab6a | D4 | siloxane |  | Cylinder | C8H25O4Si4+;297.08 | C8H27O5Si4+;315.09 |  |  |  |  | C7H21O4Si4+;281.05 | [5 empty] | 0.85 | 0.01 |  |  |  |  | 0.14
Lab6a | 1,3,5-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 | C9H11+;119.09 |  | C8H9+;105.07 | [5 empty] | 0.72 |  |  | 0.21 | 0.01 |  | 0.06
Lab6a | limonene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H11+;95.086 | C6H7+;79.054 |  |  |  | 0.41 | [5 empty] | 0.49 | 0.07 | 0.03
Lab6a | ethanol | alcohol |  | Cylinder | C2H7O+;47.049 | C2H9O2+;65.06 |  |  | C2H5O+;45.033 | [7 empty] | 0.44 | 0.05 |  |  | 0.51
Lab6a | acrolein | aldehyde | alkene | Cylinder | C3H5O+;57.033 |  |  |  | C3H3O+;55.018 | [7 empty] | 0.95 |  |  |  | 0.05
Lab6a | dimethylsulfide | sulfide |  | Cylinder | C2H7S+;63.026 | [5 empty] | CH3S+;46.995 | [5 empty] | 0.91 | [5 empty] | 0.09
Lab6a | methyl vinyl ketone | ketone | alkene | Cylinder | C4H7O+;71.049 | [5 empty] | C2H3O+;43.018 | C3H3O+;55.018 |  |  |  |  | 0.84 | [5 empty] | 0.08 | 0.08
Lab6a | cyclohexene | cyclic | alkene | Cylinder | C6H11+;83.086 | [5 empty] | C4H7+;55.054 | C5H7+;67.054 |  |  |  |  | 0.6 | [5 empty] | 0.19 | 0.21
Lab6a | D3 | siloxane |  | Cylinder | C6H19O3Si3+;223.06 | C6H21O4Si3+;243.05 | [10 empty] | 0.91 | 0.09
Lab6a | 2-hexanone | ketone |  | Cylinder | C6H13O+;101.1 | C6H15O2+;119.11 |  |  |  | C6H11+;83.086 | C4H7+;55.054 | C2H5O+;45.033 | C3H6O+;58.041 |  |  |  | 0.73 | 0.03 |  |  |  | 0.09 | 0.05 | 0.05 | 0.05
Lab6a | furfural | aromatic | aldehyde | Cylinder | C5H5O2+;97.028 |  |  |  | C5H3O2+;95.013 |  | C4H5O+;69.033 | [5 empty] | 0.87 |  |  |  | 0.1 |  | 0.03
Lab6a | styrene | aromatic | alkene | Cylinder | C8H9+;105.07 | [5 empty] | C6H7+;79.054 | [5 empty] | 0.99 | [5 empty] | 0.01
Lab6a | b-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 |  |  | C10H16+;136.12 |  |  | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 | C7H7+;91.054 |  | 0.37 |  |  | 0.02 |  |  | 0.39 | 0.13 | 0.03 | 0.03 | 0.03
Lab6a | benzaldehdye | aromatic | aldehyde | Cylinder | C7H7O+;107.05 |  |  |  | C7H5O+;105.03 |  | C6H7+;79.054 | [5 empty] | 0.7 |  |  |  | 0.18 |  | 0.12
Lab6a | o-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 |  |  | C7H7+;91.054 | [5 empty] | 0.71 |  |  | 0.17 |  |  | 0.12
Lab6a | tetrahydrofuran | heterocycle |  | Cylinder | C4H9O+;73.065 | C4H11O2+;91.075 |  |  | C4H7O+;71.049 | C4H7+;55.054 | C3H3+;39.023 | [5 empty] | 0.51 | 0.01 |  |  | 0.11 | 0.35 | 0.02
Lab6a | 1,4-dioxane | heterocycle |  | Cylinder | C4H9O2+;89.06 | C4H11O3+;107.07 |  |  | C4H7O2+;87.044 |  | C2H5O+;45.033 | [5 empty] | 0.65 | 0.05 |  |  | 0.02 |  | 0.28
Lab6a | bromodichloromethane | halogen |  | Cylinder | [6 empty] | CHBrCl+;126.89 | CHCl2+;82.945 | [10 empty] | 0.35 | 0.65
Lab6a | methyl isobutyl ketone | ketone |  | Cylinder | C6H13O+;101.1 | C6H15O2+;119.11 |  |  |  |  | C5H9O+;85.065 | C4H9+;57.07 | C3H7O+;59.049 |  |  |  | 0.64 | 0.03 |  |  |  |  | 0.06 | 0.21 | 0.06
Lab6a | dibromochloromethane | halogen |  | Cylinder | [6 empty] | CHBr2+;170.84 | CHBrCl+;126.89 | [10 empty] | 0.16 | 0.84
Lab6a | 4-ethyltoluene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 |  |  | C8H9+;105.07 | C7H9+;93.07 |  |  |  |  | 0.56 |  |  |  | 0.09 |  | 0.17 | 0.18
Lab6b | isoprene | alkene |  | Cylinder | C5H9+;69.07 |  |  | C5H8+;68.062 | C5H7+;67.054 |  | C3H5+;41.039 | C3H3+;39.023 |  |  |  |  | 0.44 |  |  | 0.11 | 0.32 |  | 0.08 | 0.05
Lab6b | acetone | ketone |  | Cylinder | C3H7O+;59.049 | [5 empty] | C2H3O+;43.018 | [5 empty] | 0.88 | [5 empty] | 0.12
Lab6b | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | [11 empty] | 1
Lab6b | acrylonitrile | nitrile | alkene | Cylinder | C3H4N+;54.034 | C3H6NO+;72.044 | [10 empty] | 0.99 | 0.01
Lab6b | 2-butanone | ketone |  | Cylinder | C4H9O+;73.065 | C4H11O2+;91.075 |  |  |  | C4H7+;55.054 | C2H3O+;43.018 | [5 empty] | 0.82 | 0.02 |  |  |  | 0.1 | 0.06
Lab6b | benzene | aromatic |  | Cylinder | C6H7+;79.054 |  |  | C6H6+;78.046 | [7 empty] | C6H7O+;95.049 | 0.49 |  |  | 0.46 | [7 empty] | 0.05
Lab6b | toluene | aromatic |  | Cylinder | C7H9+;93.07 |  |  | C7H8+;92.062 | C7H7+;91.054 | [7 empty] | 0.48 |  |  | 0.33 | 0.19
Lab6b | m-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 | C8H9+;105.07 |  | C7H7+;91.054 | [5 empty] | 0.53 |  |  | 0.28 | 0.03 |  | 0.16
Lab6b | a-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H9+;93.07 | C6H7+;79.054 |  |  |  | 0.34 | [5 empty] | 0.37 | 0.26 | 0.03
Lab6b | 1,2,4-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 | C9H11+;119.09 |  | C8H9+;105.07 | [5 empty] | 0.56 |  |  | 0.31 | 0.01 |  | 0.12
Lab6b | furan | heterocycle | alkene | Cylinder | C4H5O+;69.033 | [11 empty] | 1
Lab6b | carbontetrachloride | halogen |  | Cylinder | [6 empty] | CCl3+;116.91 | [11 empty] | 1
Lab6b | t-Amyl Ethyl Ether | ether |  | Cylinder | [6 empty] | C5H11O+;87.08 | C5H11+;71.086 | C4H9O+;73.065 | C3H7O+;59.049 | C3H7+;43.054 | C6H13O+;101.1 | [6 empty] | 0.14 | 0.27 | 0.08 | 0.35 | 0.11 | 0.05
Lab6b | ethylbenzene | aromatic |  | Cylinder | C8H11+;107.09 |  |  | C8H10+;106.08 |  |  | C7H7+;91.054 | C6H7+;79.054 |  |  |  | C6H7O+;95.049 | 0.24 |  |  | 0.11 |  |  | 0.36 | 0.27 |  |  |  | 0.02
Lab6b | D4 | siloxane |  | Cylinder | C8H25O4Si4+;297.08 | C8H27O5Si4+;315.09 |  |  |  |  | C7H21O4Si4+;281.05 | [5 empty] | 0.84 | 0.02 |  |  |  |  | 0.14
Lab6b | 1,3,5-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 | C9H11+;119.09 |  | C8H9+;105.07 | [5 empty] | 0.54 |  |  | 0.33 | 0.01 |  | 0.12
Lab6b | limonene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H11+;95.086 | C6H7+;79.054 |  |  |  | 0.38 | [5 empty] | 0.48 | 0.09 | 0.05
Lab6b | ethanol | alcohol |  | Cylinder | C2H7O+;47.049 |  |  |  | C2H5O+;45.033 | [7 empty] | 0.32 |  |  |  | 0.68
Lab6b | methyl vinyl ketone | ketone | alkene | Cylinder | C4H7O+;71.049 | [5 empty] | C3H3O+;55.018 | [5 empty] | 0.82 | [5 empty] | 0.18
Lab6b | D3 | siloxane |  | Cylinder | C6H19O3Si3+;223.06 | C6H21O4Si3+;243.05 | [10 empty] | 0.75 | 0.25
Lab6b | furfural | aromatic | aldehyde | Cylinder | C5H5O2+;97.028 |  |  |  | C5H3O2+;95.013 |  | C4H5O+;69.033 | [5 empty] | 0.73 |  |  |  | 0.24 |  | 0.03
Lab6b | styrene | aromatic | alkene | Cylinder | C8H9+;105.07 | [5 empty] | C6H7+;79.054 | [5 empty] | 0.99 | [5 empty] | 0.01
Lab6b | b-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 |  |  | C10H16+;136.12 |  |  | C6H9+;81.07 | C7H9+;93.07 | C7H11+;95.086 | C6H7+;79.054 | C7H7+;91.054 |  | 0.29 |  |  | 0.04 |  |  | 0.33 | 0.22 | 0.03 | 0.04 | 0.05
Lab6b | ethyl acetate | ester |  | Cylinder | C4H9O2+;89.06 | [5 empty] | C2H3O+;43.018 | C2H5O2+;61.028 |  |  |  |  | 0.37 | [5 empty] | 0.15 | 0.48
Lab6b | benzaldehdye | aromatic | aldehyde | Cylinder | C7H7O+;107.05 |  |  |  | C7H5O+;105.03 |  | C6H7+;79.054 | [5 empty] | 0.53 |  |  |  | 0.34 |  | 0.13
Lab7a | acetaldehyde | aldehyde |  | Cylinder | C2H5O+;45.033 | C2H7O2+;63.044 |  |  | C2H5O+;45.033 | [7 empty] | 0.75 | 0.24 |  |  | 0.01
Lab7a | limonene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H11+;95.086 | C6H7+;79.054 |  |  |  | 0.42 | [5 empty] | 0.49 | 0.07 | 0.02
Lab7a | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.62 | 0.38
Lab7a | ethanol | alcohol |  | Cylinder | C2H7O+;47.049 | C2H9O2+;65.06 | C2H11O3+;83.07 | [9 empty] | 0.4 | 0.57 | 0.03
Lab7a | acrolein | aldehyde | alkene | Cylinder | C3H5O+;57.033 | C3H7O2+;75.044 | [10 empty] | 0.9 | 0.1
Lab7a | furan | heterocycle | alkene | Cylinder | C4H5O+;69.033 | [11 empty] | 1
Lab7a | isoprene | alkene |  | Cylinder | C5H9+;69.07 | [11 empty] | 1
Lab7a | 1,4-dioxane | heterocycle |  | Cylinder | C4H9O2+;89.06 | C4H11O3+;107.07 |  |  | C4H7O2+;87.044 |  | C2H5O+;45.033 | [5 empty] | 0.57 | 0.36 |  |  | 0.01 |  | 0.06
Lab7a | toluene | aromatic |  | Cylinder | C7H9+;93.07 |  |  | C7H8+;92.062 | C7H7+;91.054 | [6 empty] | C6H7O+;95.049 | 0.65 |  |  | 0.02 | 0.2 |  | 0.13
Lab7a | 2-hexanone | ketone |  | Cylinder | C6H13O+;101.1 | C6H15O2+;119.11 | [10 empty] | 0.85 | 0.15
Lab7a | hexanal | aldehyde |  | Cylinder | C6H13O+;101.1 | C6H15O2+;119.11 |  |  |  | C6H11+;83.086 | [6 empty] | 0.19 | 0.21 |  |  |  | 0.6
Lab7a | D4 | siloxane |  | Cylinder | C8H25O4Si4+;297.08 | C8H27O5Si4+;315.09 | [10 empty] | 0.95 | 0.05
Lab7a | acetone | ketone |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 | [10 empty] | 0.96 | 0.04
Lab7b | acetaldehyde | aldehyde |  | Cylinder | C2H5O+;45.033 | C2H7O2+;63.044 | [10 empty] | 0.86 | 0.14
Lab7b | limonene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C6H7+;79.054 |  |  |  |  | 0.5 | [5 empty] | 0.48 | 0.02
Lab7b | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.74 | 0.26
Lab7b | acetone | ketone |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 | [10 empty] | 0.99 | 0.01
Lab7b | acrylonitrile | nitrile | alkene | Cylinder | C3H4N+;54.034 | C3H6NO+;72.044 | [10 empty] | 0.94 | 0.06
Lab7b | isoprene | alkene |  | Cylinder | C5H9+;69.07 |  |  | C5H8+;68.062 | C5H7+;67.054 |  | C3H3+;39.023 | C3H5+;41.039 |  |  |  |  | 0.92 |  |  | 0.01 | 0.02 |  | 0.02 | 0.03
Lab7b | methacrolein | aldehyde | alkene | Cylinder | C4H7O+;71.049 | C4H9O2+;89.06 | [10 empty] | 0.98 | 0.02
Lab7b | 2-butanone | ketone |  | Cylinder | C4H9O+;73.065 | C4H11O2+;91.075 |  |  |  | C4H7+;55.054 | [6 empty] | 0.94 | 0.01 |  |  |  | 0.05
Lab7b | benzene | aromatic |  | Cylinder | C6H7+;79.054 |  |  | C6H6+;78.046 | [7 empty] | C6H7O+;95.049 | 0.81 |  |  | 0.01 | [7 empty] | 0.18
Lab7b | toluene | aromatic |  | Cylinder | C7H9+;93.07 |  |  |  | C7H7+;91.054 | [6 empty] | C6H7O+;95.049 | 0.76 |  |  |  | 0.13 | [6 empty] | 0.11
Lab7b | chlorobenzene | aromatic | halogen | Cylinder | C6ClH6+;113.02 |  |  | C6ClH5+;112.01 | [7 empty] | C6H7O+;95.049 | 0.71 |  |  | 0.01 | [7 empty] | 0.28
Lab7b | o-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  |  | C8H9+;105.07 |  | C7H7+;91.054 | [5 empty] | 0.84 |  |  |  | 0.03 |  | 0.13
Lab7b | 1,3,5-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  | C9H12+;120.09 | C9H11+;119.09 |  | C8H9+;105.07 | [5 empty] | 0.92 |  |  | 0.01 | 0.01 |  | 0.06
Lab7b | a-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H11+;95.086 |  |  |  |  | 0.5 | [5 empty] | 0.47 | 0.03
Lab7c | acetaldehyde | aldehyde |  | Cylinder | C2H5O+;45.033 | C2H7O2+;63.044 | [10 empty] | 0.96 | 0.04
Lab7c | a-pinene | cyclic | alkene | Cylinder | C10H17+;137.13 | [5 empty] | C6H9+;81.07 | C7H11+;95.086 | C8H11+;107.09 |  |  |  | 0.47 | [5 empty] | 0.43 | 0.04 | 0.06
Lab7c | acetonitrile | nitrile |  | Cylinder | C2H4N+;42.034 | C2H6NO+;60.044 | [10 empty] | 0.93 | 0.07
Lab7c | acetone | ketone |  | Cylinder | C3H7O+;59.049 | C3H9O2+;77.06 | [10 empty] | 0.98 | 0.02
Lab7c | acrylonitrile | nitrile | alkene | Cylinder | C3H4N+;54.034 | C3H6NO+;72.044 | [10 empty] | 0.97 | 0.03
Lab7c | isoprene | alkene |  | Cylinder | C5H9+;69.07 |  |  | C5H8+;68.062 | C5H7+;67.054 |  | C3H3+;39.023 | C3H5+;41.039 |  |  |  |  | 0.92 |  |  | 0.02 | 0.02 |  | 0.01 | 0.03
Lab7c | methacrolein | aldehyde | alkene | Cylinder | C4H7O+;71.049 | C4H9O2+;89.06 | [10 empty] | 0.93 | 0.07
Lab7c | 2-butanone | ketone |  | Cylinder | C4H9O+;73.065 | C4H11O2+;91.075 | [10 empty] | 0.93 | 0.07
Lab7c | benzene | aromatic |  | Cylinder | C6H7+;79.054 |  |  | C6H6+;78.046 | [7 empty] | C6H7O+;95.049 | 0.81 |  |  | 0.02 | [7 empty] | 0.17
Lab7c | toluene | aromatic |  | Cylinder | C7H9+;93.07 |  |  |  | C7H7+;91.054 | [6 empty] | C6H7O+;95.049 | 0.72 |  |  |  | 0.17 | [6 empty] | 0.11
Lab7c | chlorobenzene | aromatic | halogen | Cylinder | C6ClH6+;113.02 |  |  | C6ClH5+;112.01 | [7 empty] | C6H7O+;95.049 | 0.75 |  |  | 0.03 | [7 empty] | 0.22
Lab7c | o-xylene | aromatic |  | Cylinder | C8H11+;107.09 |  |  |  | C8H9+;105.07 |  | C7H7+;91.054 | [5 empty] | 0.84 |  |  |  | 0.04 |  | 0.12
Lab7c | 1,3,5-trimethylbenzene | aromatic |  | Cylinder | C9H13+;121.1 |  |  |  | C9H11+;119.09 |  | C8H9+;105.07 | [5 empty] | 0.92 |  |  |  | 0.02 |  | 0.06

## 2_Lab_ID
ID | IMR Temperature (°C) | IMR Pressure (mbar) | IMR Front (V) | IMR Back (V) | IMR Front - Back (V) | E/N (Td) | IMR RF Voltage (V) | BSQ RF Voltage (V) | Skimmer (V) | BSQ Front (V) | Skimmer - IMR Back (V) | BSQ Front - Skimmer (V) | Water Flowrate (sccm) | Inlet flowrate (cm3 min-1) | Ion source voltage (V) | Date Acquired
Lab1a | 60 | 2 | 600 | 20 | 580 | 133 | 450 | 350 | -4 | -11.7 | -24 | -7.7 | 20 | 120 | 427 | 2023-05-01
Lab1b | 60 | 2 | 600 | 20 | 580 | 133 | 450 | 300 | -4 | -11.7 | -24 | -7.7 | 20 | 120 | 432 | 2024-05-01
Lab2a | 60 | 2.4 | 600 | 25 | 575 | 110 | 450 | 300 | -4 | -11.7 | -29 | -7.7 | 19 | 100 | 424 | 2020-10-01
Lab2b | 60 | 2.4 | 650 | -5 | 655 | 124 | 400 | 400 | -9.4 | -17.5 | -4.4 | -8.1 | 29 | 100 | 428 | 2023-11-01
Lab3a | 100 | 1.5 | 400 | 35 | 365 | 125 | 400 | 215 | -4.7 | -9.3 | -39.7 | -4.6 | 20 | 96 | 431 | 2020-12-01
Lab3b | 100 | 1.5 | 410 | 25 | 385 | 132 | 400 | 215 | -7 | -11 | -32 | -4 | 15 | 88 | 438 | 2022-11-01
Lab4a | 100 | 2.5 | 630 | 36 | 594 | 122 | 450 | 320 | -4.5 | -9.5 | -40.5 | -5 | 20 | 79 | 425 | 2023-12-01
Lab4b | 100 | 2.5 | 630 | 36 | 594 | 122 | 450 | 320 | -4.5 | -9.5 | -40.5 | -5 | 20 | 95 | 426 | 2024-09-01
Lab5 | 110 | 2.5 | 650 | 26 | 624 | 132 | 400 | 250 | -1 | -4.5 | -27 | -3.5 | 21 | 180 | 434 | 2021-07-01
Lab6a | 90 | 1.5 | 500 | 20 | 480 | 160 | 400 | 255 | 0.9 | -5.6 | -19.1 | -6.5 | 15 | 260 | 421 | 2021-05-01
Lab6b | 90 | 1.5 | 500 | 20 | 480 | 160 | 400 | 255 | 0.9 | -5.6 | -19.1 | -6.5 | 15 | 290 | 428 | 2022-05-01
Lab7a | 100 | 2.2 | 600 | 30 | 570 | 133 | 450 | 325 | -8.7 | -12.9 | -38.7 | -4.2 | 20 | 100 | 422 | 2022-04-01
Lab7b | 100 | 2.2 | 600 | 30 | 570 | 133 | 450 | 325 | -8.7 | -12.9 | -38.7 | -4.2 | 20 | 100 | 425 | 2022-09-01
Lab7c | 100 | 2.2 | 600 | 30 | 570 | 133 | 450 | 325 | -8.7 | -12.9 | -38.7 | -4.2 | 15 | 100 | 426 | 2023-05-01

## 3_EN_Calculator
E/N calculator |  | Enter Green Cells
Calculation in Red Cells
drift tube properties |  |  | Notes
length | 0.1 | m
dV | 580 | V | IMR Front - Back
E | 5800 | V/m | Electric field
Pressure and temperature
p | 2 | mbar
T | 100 | degrees
T | 373 | k
R | 0.083 | m3 mbar K-1 mol-1
Nv | 3.891e+22 | molecules m-3
Nv | 38909525501469680 | molecules cm-3
Td | 1e-21 |  | electric field of approximately 2400 V m2 divided by loschmidt constant
EN calculation
E/N | 1.4906e-19
E/N | 149.0638 | Td