## Worksheet 1
Test number | Test description | Instrument serial number | Instrument version | Analysis
Test 1: Multi-hour cylinder test with Nafion | 12 h cylinder tests with Nafion | A665, A792, A800, A886 | Version 2 | Water vapor uncertainty: Determine water vapor correction if possible and estimate component of uncertainty, Uw.
Allan-Werle deviation: Use time period with low and stable water vapor calculate Allan deviation to (1) compare performance between instruments, (2) determine optimal averaging time, and (3) inform needed frequency for calibration.
Test 2: Multi-hour cylinder test | 2 h cylinder test without Nafion | A112, A752, A778, A792 (50 h) | Version 2 | Uncertainty due to instrument noise: Determine needed frequency for calibration. Estimate component of uncertainty due to drift between calibration cycles, Un.
Test 3: Multi-cylinder testing | Cycle through 3 NOAA tertiary tanks and each of 17 cylinders treated as unknown, repeated | A778 | Version 1 | Cylinder calibration: Calibrate the unknown cylinders for C2H6. Quantify repeatability as standard deviation of the calibration cycles.
| 16–32 times for each unknown cylinder. CH4 is known for all tanks based on calibrations with Picarro instrument. C2H6 is known for 6 of the cylinders from NOAA. |  |  | Determine C2H6 cylinder assignment uncertainty: Estimate component of C2H6 uncertainty, Uc, due to cylinder assignment errors by comparing to NOAA values.
Test 4: Pre-deployment test | Pre-deployment test sampling room air with two calibration cylinders for each system. Concurrent Picarro CH4 measurements. Drop power for 10 s and for 12 h. | A665, A792, A800, A886 | Version 1 (8 h) and Version 2 (10 d) | Pre-deployment test: Calculate mean bias and standard deviation of hourly calibrated values for CH4 and C2H6. Compare to total uncertainty, Ut. Test calibration approach using both high and ambient cylinders every hour compared using high cylinder every 25 h and ambient cylinder every hour.
Warm-up test: Determine necessary length of warm-up time for field measurements by examining calibration gas response.
Test 5: Field test | Co-located at NOAA CAO tower (CH4). | A778, A665 | Version 1 and Version 2 | Field test: Calculate mean bias and standard deviation of hourly calibrated values for CH4. Compare to total uncertainty, Ut.