Interactive comment on “Response of the Nevzorov hot wire probe in Arctic clouds dominated by very large droplet sizes”

A. Schwarzenboeck et al.

Referee #1: Walter Strapp

General Comments:

This article contains valuable new information on the Nevzorov LWC/TWC probe and

should be published. I am glad to see someone is still doing this kind of work, as it is

essential to do detailed instrument studies to make progress in cloud physics. The article

does contain some errors that need to be corrected. I also think it would be a much better

article if authors include some additional discussion of the uncertainties that I have noted

in the “Specific Comments” below. I have also made some recommendations throughout

the text that the authors may choose at their discretion to implement.

First and foremost we would like to thank the reviewer for his valuable comments and ideas to improve the quality of the manuscript. Please find below a detailed item-by-item response, discussing the reviewer’s remarks. The manuscript has been improved with respect to the below discussed items.

Specific Comments:

Page 1294, line 22: Is the LWC-100 the DMT version of the King Probe? If so, please

add DMT to the sentence.

DMT has been added to the sentence.

Page 1295, line 1: After the end of the sentence (beyond 40 μm), I recommend that you

refer to Biter et al. (1987) and Strapp et al (2003). This evaporation and breakup effect is

not a theoretical result; it is rather an empirical observation first measured by Biter.

Biter C.J., Dye, J.E. Huffman, D. and King, W.D. 1987 The drop response of the CSIRO liquid water

content. J. Atmos. Ocean Techn., 4, 359-367

The text has been modified and the above mentioned reference added.

Page 1295, line 3: The PVM also suffers from rolloff with MVD. You might want to

refer to Manfred Wendish’s paper.

Wendisch, M., T.J. Garrett, and J.W. Strapp, 2002: Wind Tunnel Tests of the Airborne PVM-100A

Response to Large Droplets. J. Atmos. Oceanic Technol., 19, 1577–1584.

The above reference has been added.

Page 1295, starting at line 6: The sentence is not quite accurate. It is easier for me to

rewrite the sentence than try to explain it. I think it will be self-explanatory. The FSSP is

not an imaging probe.

“Besides these direct measurements of LWC, estimates of the total condensed water

content, including the ice phase (IWC and/or LWC) can be derived by integrating the size

spectra from the Forward Scattering Spectrometer Probe FSSP (Baumgardner, 1985), and

the spectra derived from imagery from the 2D-C (Knollenberg, 1981) or the Cloud

Particle Imager CPI (Lawson, 1998).”

The text suggested by the reviewer has been adopted.

Page 1295, line 14: Recommend change to “The CVI technology has been recently

integrated into an oversize PMS canister, including hygrometer, by Droplet Measurement

Technologies (DMT).”

The improved text suggested by the reviewer has been adopted.

Page 1296, line 8: I don’t understand the sentence starting with “ The quite good

correlation ..”. This seems to imply that the probe fails in large droplets or the ice phase.

In the case of the ice phase, isn’t this the principle of operation of the phase

discrimination, that the TWC will read higher in ice than in liquid? I would suggest

changing the sentence to the following:

“The Nevzorov LWC and TWC sensors correlate very well in small droplet cloud. In

mixed phase clouds, the TWC sensor measurement exceeds the LWC measurement due

to its much higher efficiency in detecting ice particles, as expected due to the

aerodynamic design of the capture volume. In liquid-only large-droplet conditions, the

TWC measurement exceeds the LWC measurement, because the TWC sensor has much

lower re-entrainment losses, and due to the fact that the existing collection efficiency

assumptions for these probes in such conditions have not been well characterized to

date.”

The improved text has been adopted.

Page 1296, line 13: This is not correct. The Strapp et al. (2003) study presents results as

a function of the median volume diameter. This is the diameter in the LWC versus

diameter distribution at which 50% of the water is below, and 50% is above. Usually, the

median volume diameter is higher than the mean volume diameter, which you have

shown on page 1303. Also, you are missing a ‘pi’ and a 6 in equation 10, although we

didn’t use that anyway.

Furthermore, the drop distributions in a wind tunnel are typically not narrow and

monodisperse as you state in the text. There are quite wide. The large MVDs are

normally produced by long tails in the distributions. Normally, the LWC-versusdiameter

forms a ‘bell’ type distribution. The biggest difference with atmospheric

distributions is the lack of a mode in the small-droplet part of the distribution. The

droplet concentrations continue to increase with decreasing diameter. This causes

problems for the FSSP, especially for the lower MVDs, because the drop concentrations

are so high. You will need to rewrite all the sections that discuss the IRT distributions. If

you would like to send it to me for inspection, I would be happy to do it.

I am really sorry for my misunderstanding. After reading your publication (Strapp et al., 2003) I’ve just looked up the definition of the median volume diameter MVD in Hind’s book, where MVD, however, is given more as a diameter of average volume  . Unfortunately, I did not search for a cross-check to confirm or reject that definition. As a consequence, I had to assume very narrow IRT droplet spectra to justify the use of an average volume diameter as the spectrum reference diameter (representative in volume) for the wind tunnel study (Strapp et al., 2003). Again, I am sorry for that misinterpretation, since it turned out to be an error. The correct interpretation of MVD - where MVD is defined as  - represents a much better means to express drop size in terms of the volume of liquid sprayed into the IRT tunnel.

Related to the above definition of MVD in our study we have been using the volume mean diameter VMD  defined as .  This volume-weighted average diameter VMD is a closely related parameter to the median volume diameter MVD used in the Strapp study (MVD is often used to describe artificially produced droplet populations).

As a consequence of the author’s error in the comprehension of MVD derived for IRT droplet spectra, in the revised version of the manuscript all the sections that discuss the IRT distributions have been rewritten.

Page 1296, line 22: Please consider adding the following sentences and references

before “Further tests”. This covers some earlier work on the high-speed video images of

ice particles bouncing from hot-wires, and the resulting underestimates of IWC from hot

wires, including the Nevzorov TWC.

“However, a series of tests at the Cox and Co. wind tunnel, using ice shaved from blocks

to simulate ice particles, revealed that a fraction of these ice particles bounce out of the

sample volumes of various hot wire devices, including the Nevzorov TWC probe,

resulting in an underestimate of the IWC measurement (Emery et al. 2004, Strapp et al.

2005). “Further tests …” and then take out the redundant information after that. If youdecide to take my suggestion and rewrite the paragraph, I would be happy to inspect the

new paragraph after you have completed it.

Emery, E., Miller, D., Plaskon, S., Strapp, J.W., and Lilie, L.E., 2004: Ice Particle Impact on Cloud Water

Content Instrumentation, 42nd AIAA Aerospace Sciences Meeting and Exhibit, Jan. 2004, AIAA-2004-

0731.

Strapp, J.W., Lilie, L.E., Emery, E.E., and Miller, D.R., 2005: Preliminary Comparison of Ice Water

Content as Measured by Hot Wire Instruments of Varying Configuration, 43rd Aerospace Sciences

Meeting and Exhibit, Reno, NV, 11-13 Jan. 2005, AIAA-2005-0860.

The above text suggested by the reviewer has been added, the subsequent text of the paragraph has been shortened.

Page 1297, line 25: What type of instrument was used to measure aerosol, and what was

its minimum size limit? This is relevant to the comment that concentrations were low.

The mentioned aerosol particle concentrations are those having been measured with a CPC 3010. Thus, this includes all particles larger than 11 nm. The CPC 3010 measurement device for total aerosol concentrations now is stated in the text.

Page 1298, line 7: Regarding “feeder-seeder”, do you mean that ice crystals from a

higher cloud were falling into the lower cloud? If yes, please state explicitly that there

were higher clouds.

The feeder-seeder phenomenon could be observed few times from lidar observations on board Polar2, during flight sequences dedicated to out of cloud measurements (lidar measurements beyond cloud layer).

Page 1301, line 20:

Please consider adding the next sentence. I think this adds value by showing that this

result is common to other hot-wire cylinders, and therefore it is not specific to the

Nevzorov LWC. In my mind, this makes the result more believable.

“Similar results were first shown by Biter et al. for the King probe, another cylindrical

hot wire like the Nevzorov LWC, and for various other cylindrical hot wires by Strapp et

al. (2003). “

The text suggested by the reviewer has been added in the revised version of the manuscript.

Page 1302, line 15: Is not the 1.595 * V**2(LWC)/V**2 (TWC) the same as

LWC(uncorrected)/TWC(uncorrected). I don’t quite understand the reason for carrying

the voltages at this stage of the article. It makes it more difficult to conceptualize.

However, it is not a big issue.

The authors’ idea of using the raw signals (voltages or power) is to discuss efficiencies before calculating condensed water contents (IWC, LWC) based on these efficiencies. Our concept has been first to derive/discuss efficiencies from raw signals and then water contents with the help of these efficiencies. We decided to maintain our concept for the revised version.

Page 1303, lines 2-7: same comment as Page 1296, line 13: This is an error.

The paragraph has been modified - see also the authors’ response given above for page 1296, line 13.

Page 1303: line 14: Regarding Figure 7, I have two major points (1) and (2), and a minor

point (3):

(1) Strapp et al. (2003) discuss with some detail that the measurement of MVD is

very difficult when you need to use more than one probe (i.e. when the MVD is

large), and this is probably the worst for MVDs in the range of 30-100 μm. In

your own paper you state: “below 150 μm the (2D-C) is underestimating the

crystal concentration”, and therefore presumably also the droplet concentration.

Your paper needs to have some kind of a discussion about the uncertainties in

MVD. For example, someone might use your correction equation using their

own MVD estimate that might be highly biased to yours, especially if they are

using different cloud probes. A discussion is necessary to warn the reader that

MVD estimation can be difficult with common probes. This could be causingsome of the difference between Strapp et al. (2003) and your results in Fig. 7. It

would be nice if this were stated in your text, so that the differences between

Strapp et al and the present article have some possible explanation (i.e. MVD).

Also, please note that your results in Fig. 7 show how to correct the Nevzorov

LWC so that it is the same as the Nevzorov TWC. This is OK, but it is maybe

not so obvious to all readers. The results of Strapp et al. display the response of

the instrument to its readings at 20 microns MVD, which is not exactly the same

thing as yours.

It is clear that it is difficult to derive an absolute estimate of the uncertainty in MVD or VMD of one particular probe (or a particular probe combination). Differences may arise from individual probe performances/characteristics of the same probe type, data processing software, and other probe and/or operator specific parameters. Already the particular shape of the cloud droplet size spectra (for example comparing narrow to broad size spectra), would make VMD or MVD a more or less suitable single reference parameter to quantify the Nevzorov probe response to an ensemble of droplet sizes. In other words, one and the same MVD might stem from different droplet size spectra, whereas the probe might react differently concerning these different size spectra having the same MVD.

Thus, in the revised version of the manuscript we have added a paragraph to discuss potential uncertainties in MVD estimates from cloud probes and in particular to warn the reader/user that MVD estimation from common cloud probes could be biased to a certain extent.

(2) How do you propose to use your equation (14) in mixed phase situations, as you

have in section 5.2? Accurate determination of the liquid MVD from particle

probes in mixed-phase conditions is exceedingly difficult. Maybe you can do it

with the CPI … I am not sure. It depends on whether you believe the CPI

distributions, and the separation of the ice particles from the water droplets.

Please explain this in the text somewhere.

The ASTAR 2004 measurement campaign produced our first CPI data set and we put a lot of effort to improve the shape classification of the CPI standard software. The CPI images were preprocessed using the software (Lefevre, 2007) developed at the Laboratoire de Météorologie Physique (LaMP). The LaMP software is based on the features and algorithms outlined in the manual of the original CPIview software (see CPIview, 2005; Lawson et al., 2001; Baker and Lawson, 2006). Moreover, it provides additional information on the ice-particle morphology that is not available from the CPIview software. More specifically, we added, to the original automated habit classification, three habit types and used our algorithm to evaluate the effective sampling volume of the CPI (Lefèvre, 2007). The output of our new software has been successfully compared to those from the original CPIview software (CPIview, 2005). In addition, the automatic classification of CPI images has been refined by human classification for the entire ASTAR dataset (Lefèvre, 2007). To be able to calculate IWC and LWC in mixed phase clouds for the Nevzorov probe, thus, using size and phase dependent Nevzorov probe efficiencies, we have to determine/calculate the phase and MVD from size distributions for liquid and solid cloud particles. The most critical point is of course the correct phase discrimination (surface roughness) for spherical particles, which is more difficult to determine for smallest spherical particles. However, we are convinced that beyond a diameter of 50 µm we can be quite sure from CPI images about the phase (liquid or ice) of spherical particles captured on CPI images. We added some text in that sense to explain potential ambiguities/uncertainties related to phase discrimination from CPI images.

(3) What is the potential effect of small gain differences between the LWC and TWC

probe in Figure 7? We cannot measure the calibration constants perfectly, so

there will be some linear effect. I suspect it is small, but it is worth a few lines in

the text somewhere.

We added a sentence to indicate the potential of uncertainties in the Nevzorov probe calibration constants.

Page 1304, line 5: same comment as Page 1296, line 13:

The sentence has been modified - see also the authors’ response given above for page 1296, line 13.

Page 1304, line 15 and beyond:

Strapp et al. 2003 showed collection efficiencies from the LEWICE CFD model as

follows:

Do the above efficiencies affect your results at all?

First of all, thank you mentioning the modelled TWC efficiency from LEWICE CFD tool. A difference in TWC efficiency proposed by the model and suggested from Korolev et al. can be seen roughly between 5 and 15 (20) µm. This is the size range where only few VMD data points (from size distributions) are available for our ASTAR dataset (data points with significant scattering!). The presented ASTAR measurements are less powerful to cover the size range of smallest droplets. Of course the recalculation of data points using the digitized curve from LEWICE model slightly changes data points in figure 8 with respect to smallest VMDs. Since we suggest our efficiency for LWC,droplets to be valid only for VMDs beyond 25 µm and in addition the proposed efficiency has been fitted to LWC,droplets proposed by Korolev (1998), the impact of ‘minor’ variations for TWC,droplets in the size range 5-20µm is of very limited effect for the results presented here. In the revised manuscript the authors added a sentence to mention the LEWICE model results in Strapp et al. 2003.

Page 1305, line 1: same comment as Page 1296, line 13:

The sentence has been modified - see also the authors’ response given above for page 1296, line 13.

Page 1306, section 5.2:

First, see comment Page 1303: line 14, item (2)

See response given above concerning phase discrimination, and evaluation of MVD from measurements.

Second, it is almost certain that the standard Nevzorov TWC is about a factor of 3 low in

situations where large ice crystals are not present. This has been the case in wind tunnels

and in natural cloud studies, when compared to both the CVI and the deep cone. To quote

Alexei Korolev’s ICCP 2008 paper,

“For ice particle spectra with Dmax<4mm, the IWC measured by the standard

Nevzorov shallow cone is approximately 3±0.2 times lower than that measured by the

Nevzorov deep cone and CSI. Assuming that the close agreement between the CSI

and the deep cone indicates that both are measuring approximately correctly, this

allows for corrections of IWC data sets collected with the Nevzorov shallow cone

TWC sensor during previous flight campaigns, in cases where large particles are not

present. The correction must take into account any liquid fraction in the cloud, which

will be measured at a much higher efficiency.”

Did you have particles larger than 4 mm in your arctic cases? I am not sure why you

don’t just use a factor of 3 correction and apply it to Fig. 9. This figure is for illustration

of a real application, and you might as well use the best estimates of the efficiencies.

As can be seen in the section 5.2 we are aware of the discussion of improved efficiency for IWC recovery from the deep TWC cone as compared to the shallow one. We considered the results presented in figure 9 as one real application of the presented study of Nevzorov probe efficiencies. This result, however, is still subject to a potential modification due to the ongoing discussion of poor IWC recovery by the shallow cone. According to our understanding, a simple correction factor of 3 may not be the final conclusion, since the collection efficiency εTWC,crystals  of the TWC shallow cone with respect to ice should be a function of diameter (VMD, MVD).

What we still don’t understand is the fact that in former days Nevzorov IWC recovery has been well correlated to other retrievals (2D-C or combination 2D-C + 2D-P).

See for example:

- Isaac et al., AIAA 2006-1221: 2D-P IWC fits really well to TWC of Nevzorov (shallow, because 3.5 mm deep cone).

- Korolev & Strapp 2002-0679: Almost perfect correlation of IWC derived from 2D-C + 2D-P and Nevzorov TWC sensor (probably the shallow cone!).

This is why we didn’t dare in the manuscript to simply apply a factor of 3 to correct IWC recovery for the shallow cone. By the way, do we then have to conclude that 2D-C and 2D-P probes are underestimating IWC by a factor of 3?

Finally we think that the reviewer’s argument is o.k., why not applying the nowadays best estimate for εTWC,crystals.

Thus, in the revised version of the manuscript we propose to use the modified efficiency for our shallow TWC cone of . Figure 9 was produced under that assumption, all the text in chapter 5.2 has been corrected accordingly.